P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Analyzer, Automated Molecular detection for infectious markers

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Ms Stella CHEUNG)

### [by fax: 2127 4795 or email: syhcheung@healthbureau.gov.hk]

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply of Analyzer, Automated Molecular detection for infectious markers (hereinafter refers as the “System”) to The Chinese Medicine Hospital of Hong Kong (“opCMHHK”) located at 1 Pak Shing Kok in Tseung Kwan O, New Territroies, Hong Kong.  The CMHPO therefore wishes to collect market information on this systems.

- Background of the CMHHK Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Analyzer, Automated Molecular detection for infectious markers that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Analyzer, Automated Molecular detection for infectious markers.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ________________________________________ (Please also provide the expected life of these excluded components) |
| - *Total weight of the proposed System | ____________kg |

*	The maximum floor loading capacity where the System is to be installed is 5 kPa. Please ensure that your proposed System can comply with this requirement.

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “System” in this Part 3 refers to section A1.1 to A1.8 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient.

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A. | Technical Requirements |  |  |
| 1. | System Features |  |  |
|  | The System offered shall be a fully automated clinical diagnostic system using Real-Time Polymerase Chain Reaction (RT-PCR) technology to perform clinical sample preparation, nucleic acid extraction, amplification and detection automatically in one single processing unit. |  |  |
|  | The System shall compose of at least 4 processing units and one set of workstation with required software to operate the system. |  |  |
|  | The System shall have random access features such that assays up to 4 different nucleic acid targets or more can be processed at the same time. To avoid batching samples into a single run, the System shall enable setup of an assay in any processing unit while other modules are running. |  |  |
|  | Sample throughput shall allow simultaneous run for at least 4 tests per run. |  |  |
|  | The System shall have feature such that the processing unit can be added up to 4 processing units. |  |  |
|  | The System shall not be required to operate in a negative pressure PCR laboratory room. |  |  |
|  | The System shall consist of a barcode scanner or functionally equivalent design for scanning the specimen’s sample ID barcode, and the cartridge barcode, which enters the cartridge type and settings into the system. |  |  |
|  | The System shall be CE-IVD marked, or FDA approved or equivalent regional medical device regulatory approval/clearance. |  |  |
|  | External dimension: 280mm (W) x 310mm (D) x 300mm (H) ± 2%. |  |  |
|  | The System offered must be suitable for use throughout the environmental ranges unless otherwise specified: Temperature 10 °C - 35 °C, relative humidity 30% - 85% or wider. |  |  |
|  | Processing Unit |  |  |
|  | The processing unit shall have at least 10-channel optics for the fluorescent detection of multiple unique targets for each single cartridge. The module can detect original – CF1, FAM, A532 (CF3), TxR (CF4), A647 (CF5), CF6, CF7, CF8, CF9 and CF10 signal. |  |  |
|  | The characteristics of the optical channels (or more) shall be as follows: Channel; Extraction (nm); Emissions (nm). 1; 375-405; 420-480. 2; 450-495; 510-535. 3; 500-550; 565-590. 4; 555-590; 606-650. 5; 630-650; 665-685. 6; 630-650; >700. 7; 450-495; 565-590. 8; 500-550; 606-650. 9; 450-495; 606-650. 10; 500-550; >700. |  |  |
|  | Thermal performance |  |  |
|  | Heating ramp rates: at least 4.4°C /sec or more from 50 °C to 95°C. |  |  |
|  | Cooling ramp rates (max.): at least 2.2 °C/sec or more from 95 °C to 50 °C. |  |  |
|  | Temperature duration accuracy: ±1.0 sec from programmed time or better. |  |  |
|  | Temperature accuracy: ± 1.0 °C from 60 °C to 95 °C or better. |  |  |
|  | Melt curve programmable ramp rates: 0.1 °C/sec to 1.0 ° C/sec. |  |  |
|  | The automated process of nucleic acid extraction, amplification and target detection shall be carried out in a self-contained, single use cartridge/kit placed into a processing module of the processing unit. |  |  |
|  | Data Management system |  |  |
|  | The system shall supply with one computer workstation with data management for controlling of the system and reporting of assay results. |  |  |
|  | The computer workstation shall have the following specifications for operation of the processing unit shall be provided. |  |  |
|  | Processor: Core i5 or equivalent, 1.6 GHz or better. |  |  |
|  | Hard disk storage: 512 GB or more |  |  |
|  | At least include 2 USB connections or more |  |  |
|  | Screen size: at least 14 inches or above LCD/LED display |  |  |
|  | Scanner: Barcode scanner shall be provided for the operation of the processing unit |  |  |
|  | Operating system: preloaded WIN 10 or later version. |  |  |
|  | User interface |  |  |
|  | The software shall be able to manage and control the operation of all offered modules. |  |  |
|  | The software shall be able to collect, store and analyze data generated by each module. |  |  |
|  | The software shall display the status of all processing runs, along with the sample ID, test type, and time remaining until completion or equivalent information. |  |  |
|  | The software shall allow operators to browse for runs, display results, print report and send results to LIS. |  |  |
|  | The software shall allow operators to view system details, messages and errors. If an error exists, operators shall be able to create an error bundle for exporting data for troubleshooting. |  |  |
|  | The software shall also allow operators to import assay protocols and export result files. |  |  |
|  | The Supplier shall bear the installation cost of middleware, if any and not more than two numbers of Medical Network Data Ports for each System. |  |  |
|  | The Supplier shall implement the requirements regarding LIS/Equipment Interface for Acquisitions upon contract award. |  |  |
|  | Reagent requirement |  |  |
|  | Test kit reagents shall be CE-IVD or FDA approved. |  |  |
|  | The Test kit reagents list shall detect the following target genes: |  |  |
|  | Respiratory |  |  |
|  | Detection and differentiation of SARS-CoV-2, influenza A, influenza B, and respiratory syncytial virus (RSV) |  |  |
|  | Detection and differentiation of influenza A, influenza B, and RSV |  |  |
|  | Healthcare-associated Infection |  |  |
|  | Detection of Methicillin-Resistant Staphylococcus aureus (MRSA) directly from nasal swabs in patients at risk for nasal colonization |  |  |
|  | Detection and differentiation of KPC, NDM, VIM, IMP-1, and OXA-48 in Carbapenemase Producing Enterobacteriaceae (CPE) |  |  |
|  | Detection of vancomycin-resistance (vanA-vanB) genes in Vancomycin-resistant Enterococci (VRE) |  |  |
|  | Tuberculosis diseases |  |  |
|  | Detection of Mycobacterium tuberculosis complex and Rifampin-resistance associated mutations |  |  |
|  | Blood Virology, Women’s Health, Sexual Health |  |  |
|  | Detection and differentiation of genomic DNA from Chlamydia trachomatis (CT) and/or Neisseria gonorrhoeae (NG) |  |  |
|  | Detection of high-risk Human Papillomavirus (HPV) - identifies types HPV 16 and HPV 18/45; reports 11 other high risk types in a pooled result |  |  |
|  | Detection of Group B Streptococcus (GBS) |  |  |
|  | Detection and quantitation of Hepatitis B Virus (HBV) |  |  |
|  | Detection and quantitation of Hepatitis C Virus (HCV) |  |  |
|  | LIS/Analyzer interface |  |  |
|  | The system shall be interface with the CMHHK Laboratory Information System (LIS) to allow result upload and download between the analyzer and LIS immediately when result is available. |  |  |
|  | The successful Supplier shall be responsible and bear all the cost, including the material, hardware, and software, resources of any kind, manpower and expert support, required for the successful LIS/Analyzer interface of each unit of analyzer. |  |  |
|  | All hardware and software setup must be configurable and modifiable in order to meet the special requirements of LIS connection, if any. |  |  |
|  | All Connection and interface should comply with the “Term and Conditions for Medical System Network connecting to the Hospital Network” which shall include the installation of Medical Network Data Port(s) for instrument connection. |  |  |
|  | Regardless of where the problems may arise, the successful Supplier must provide all the necessary assistance, consultancy, time and resources, including any hardware, as requested by the CMHHK to ensure the successful completion of the LIS/Analyser interface. |  |  |
|  | Should there be operational reasons that the system resided within the CMHHK premises needs to connect to the CMHHK network for information exchange with the CMHHK IT systems, the successful Supplier shall be responsible for implementing such network connection and ensure that all requirements described in “IT Security Requirements for non-IT system” are strictly adhere to. |  |  |
|  | Other Requirements |  |  |
|  | Equipment offered shall comply with the safety requirements of IEC60601-1 / IEC61010-1 or equivalent. |  |  |
|  | Equipment offered shall comply with the electromagnetic compatibility (EMC) requirements of IEC60601-1-2 / IEC61326-1 or equivalent. |  |  |
|  | The System offered shall comply with the Particular requirements for in vitro diagnostic (IVD) medical equipment of IEC61010-2-101 or equivalent standard. |  |  |
|  | The System offered shall comply with the IVD equipment electromagnetic compatibility (EMC) requirements of IEC61326-2-6 or equivalent |  |  |
|  | The uninterrupted power supply (UPS) offered shall comply with IEC62040-1, IEC62040-2 and IEC62040-3, or equivalent standards |  |  |
|  | The equipment shall remain operational and within specification throughout the voltage range of 220V ± 6%, 50Hz ± 2%, 1-phase A.C. electrical supply. |  |  |
|  | Single-phase mains operated equipment shall be fitted with a power plug suitable for the site condition. The plug shall comply with relevant standards e.g. BS1363 for 13A plug if applicable. |  |  |
|  | The System shall be provided with an over-current protection cutout device. The System shall be effectively bonded to earth unless it is double insulated. |  |  |
|  | All components shall be free of burrs, sharp edges, protrusion and other defects which may cause hazard to operator. All surfaces and edges shall be smooth and non-abrasive. |  |  |
|  | The requirements for flexible power cords of the System including the marking requirements on the outer sheath, the requirements on colour identification, conductors and insulation shall comply with the relevant standards. For example, the thickness of the insulation shall not be less than the requirements specified in IEC60227 or IEC60245 or EN 50525 or equivalent. |  |  |
|  | The System shall be connected to uninterruptible Power Supply (UPS) at no additional charge and the UPS shall support for at least 15 minutes or at least one complete IVD analysis run cycle (whichever is longer) of power dripping and assessable for maintenance. |  |  |
|  | The UPS system shall provide at least visual indication for overload, over-temperature, critical breaker open, low battery voltage and input power failure. |  |  |
|  | The Supplier shall provide at least one year guarantee for the System supplied, or any part or portion thereof, starting from the date of acceptance of the System. During the guarantee period all services, which include replacement of faulty parts, scheduled and breakdown services by qualified maintenance personnel, shall be provided free of charge. |  |  |
|  | Implementation Services |  |  |
|  | The System shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included within 10 weeks from the date informed by the CMHHK. |  |  |
|  | The Supplier shall be responsible for connecting all water inlets, waste outlets and electricity supply to the System and any modification works if required. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier should note that they will be held responsible for any damage to the CMHHK property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
|  | The Supplier shall be responsible for the make good of ceiling, wall and floor in the installation site after completion of equipment installation. |  |  |
| C | Training |  |  |
|  | On-site maintenance and operational training shall be provided at no additional charges for a minimum of two maintenance and operation staff. |  |  |
|  | The Supplier shall be responsible to provide session of on-site maintenance training to representatives of CMHHK upon request. The course shall cover basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures , calibration and alignment, adjustment, etc. |  |  |
|  | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. Detailed syllabuses shall be submitted for approval, upon request. The practical part of the training shall coincide with the installation and commissioning of the System. |  |  |
|  | Training protocol of operator and content shall be submitted together with the tender return for evaluation Multiple operator training shall be required upon the request of laboratory staff and certificate shall be provided after the training as the requirement of laboratory accreditation. |  |  |
| D | Documentation |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |
|  | At the time of delivery of the equipment, appropriate Set(s) of the manufacturer’s original operation and maintenance manuals in English or in Chinese complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to the CMHHK after order placement but before the equipment delivery for documentation purpose. |  |  |
|  | The Successful Supplier shall provide the authorization letter from manufacturer for providing products, parts and service to end user. The written undertaking shall be signed by a duly authorized representative of the manufacturer and dated no later than the quotation closing date. |  |  |
| E | Acceptance Tests |  |  |
| 1. | Safety Test For the purpose of this contract the System shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Authority Representative or his authorized nominee or nominees. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the System. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
| 2. | Functional Test For the purpose of this Contract the System shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. In the event that the equipment fails to conform to the above stated requirements, the successful Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the System shall be determined by the CMHHK based upon the satisfactory completion of such functional test. |  |  |
| 3. | The successful Supplier shall provide all testing instruments, reagents and consumables to conduct site acceptance tests. All testing instruments to be used for the acceptance test shall be calibrated and copies of calibration certificates or other supporting documents shall be forwarded to the representatives of the CMHHK for records. |  |  |
| F | Indicative Warranty Service |  |  |
|  | The Supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
|  | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least two times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |
|  | The Supplier shall be responsible to make good to the satisfaction of the CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure which may arise within the warranty period of the equipment. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Analyzer, Automated Molecular detection for infectious markers can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Analyzer, Automated Molecular detection for infectious markers does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Analyzer, Automated Molecular detection for infectious markers in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| ## IEC60601-1 or equivalent / IEC61010-1 or equivalent | ## Details please refer to the Section 5.1 in Part 3 |  |  |  |
| IEC60601-1-2 or equivalent / IEC61326-1 or equivalent | Details please refer to the Section 5.2 in Part 3 |  |  |  |
| IEC61010-2-101 or equivalent standard | Details please refer to the Section 5.3 in Part 3 |  |  |  |
| IEC61326-2-6 or equivalent | Details please refer to the Section 5.4 in Part 3 |  |  |  |
| IEC62040-1 or equivalent standard, IEC62040-2 or equivalent standard, and IEC62040-3 or equivalent standard | Details please refer to the Section 5.5 in Part 3 |  |  |  |
| BS1363 13A standard or equivalent | Details please refer to the Section 5.7 in Part 3 |  |  |  |
| IEC60227 or equivalent / IEC60245 or equivalent / EN50525 or equivalent | Details please refer to the Section 5.10 in Part 3 |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Analyzer, Automated Molecular detection for infectious markers, as more particularly specified in section A1.1 to A1.8 in Part 3, including the provision of a minimum 12-month warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-month period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMHHK Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END