P R O F O R M A

Request for Market Information (“RFI”) for

Supply and Installation of Analyzers, Point-Of-Care, Whole Blood, Gas/Ph/Electrolyte

to the Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Mr Siu Lun CHENG)

### [by fax: 2127 4795 or email: slcheng@healthbureau.gov.hk]

Your ref:  --

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 8 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and installation of Analyzers, Point-Of-Care, Whole Blood, Gas/Ph/Electrolyte (hereinafter refers as the “System”) for the Chinese Medicine Hospital (“opCMHHK”) located at Pak Shing Kok in Tseung Kwan O.  The CMHPO therefore wishes to collect market information on computed tomography systems.

- Background of the CMHHK Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Analyzers, Point-Of-Care, Whole Blood, Gas/Ph/Electrolyte that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Analyzers, Point-Of-Care, Whole Blood, Gas/Ph/Electrolyte.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section G in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ________________________________________ (Please also provide the expected life of these excluded components) |
| - *Total weight of the proposed System | ____________kg |

*	The maximum floor loading capacity where the System is to be installed is 5 kPa. Please ensure that your proposed System can comply with this requirement.

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “System” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II |  |  | Column III | Column IV |
| --- | --- | --- | --- | --- | --- |
| Section | Technical Specification |  |  | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  |  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |  |  |
|  | General Description |  |  |  |  |
|  | Two (2) units of whole blood analyzers for testing of blood gas, pH, electrolytes, and other parameters are called for supply and installation, as the specifications are indicated below. |  |  |  |  |
|  | Main System |  |  |  |  |
|  | The analyzer shall adopt self-contained disposable reagent cartridges, and no extraneous equipment such as gas tanks and electrodes shall be required for calibration and operation. |  |  |  |  |
|  | The fluidic pathway should be within the self-contained cartridge from the sample probe to the waste bag, not flowing externally from the cartridge. |  |  |  |  |
|  | The analyzer shall adopt various specimen types including whole blood, capillary, mixed venous, arterial and venous blood samples. |  |  |  |  |
|  | The analyzer shall directly measure the following parameters in heparinized whole blood: |  |  |  |  |
|  | - pH, pCO2, and pO2 |  |  |  |  |
|  | - Sodium, Potassium, Ionized Calcium, Chloride |  |  |  |  |
|  | - Glucose |  |  |  |  |
|  | - Lactate |  |  |  |  |
|  | - Hematocrit |  |  |  |  |
|  | - tHb, O2Hb, COHb, MetHb, HHb |  |  |  |  |
|  | - Total bilirubin |  |  |  |  |
|  | The following derived parameters shall be calculated automatically after each measurement: |  |  |  |  |
|  | (a) Actual and standard bicarbonate, Total carbon dioxide |  |  |  |  |
|  | (b) Base excess of blood and of extracellular fluid |  |  |  |  |
|  | (c) Estimated oxygen saturation |  |  |  |  |
|  | (d) Ionized calcium standardized to pH 7.4 |  |  |  |  |
|  | (e) Anion Gap |  |  |  |  |
|  | (f) Total hemoglobin and Hematocrit |  |  |  |  |
|  | The measurement methodology shall be: |  |  |  |  |
|  | (a) Amperometric: pO2, glucose, lactate |  |  |  |  |
|  | (b) Potentiometric: pH, pCO2, Na+, K+, Ca++, Cl- |  |  |  |  |
|  | (c) Conductivity: Haematocrit |  |  |  |  |
|  | (d) Optical measurement following chemical lysing of the whole blood sample: tHb, O2Hb, COHb, MetHb, HHb and total bilirubin |  |  |  |  |
|  | The characteristics of blood gas and electrolyte measurement shall comply with the following specifications or better: |  |  |  |  |
|  | Parameter | Reportable Range | Resolution |  |  |
|  | (a) pH | 7.00 - 7.92 | 0.01 |  |  |
|  | (b) pCO2 | 6-125 mmHg (0.8-16.7 kPa) | 1 (0.1) |  |  |
|  | (c) pO2 | 6-690 mmHg (0.8-92 kPa ) | 1 (0.1) |  |  |
|  | (d) Na+ | 100-180 mmol/L | 1 |  |  |
|  | (e) K+ | 1.0-19.0 mmol/L | 0.1 |  |  |
|  | (f) Ca++ | 0.11 to 4.25 mmol/L | 0.01 |  |  |
|  | (g) Cl- | 40-158 mmol/L | 1 |  |  |
|  | (h) Glu | 4-685 mg/dL | 1 |  |  |
|  | (i) Lac | 0.3-17.0 mmol/L | 0.1 |  |  |
|  | (j) tHb | 3.0-23.0 g/dL | 0.1 |  |  |
|  | (k) O2Hb | 0.0-100% | 0.1 |  |  |
|  | (l) COHb | 0.0-75% | 0.1 |  |  |
|  | (m) MetHb | 0.0-30% | 0.1 |  |  |
|  | (n) HHb | 0.0-100% | 0.1 |  |  |
|  | (o) tBili | 2.0-40.0 mg/dL (34-684 µmol/L) | 0.1 (1) |  |  |
|  | The analyzer shall show only one single on-board expiration date on screen applied to all components including all the sensors, solutions, sample probe and waste bag to reduce user time on stock maintenance. |  |  |  |  |
|  | Patient data, QC, and quality management process can be archived and exported onto a Compact Disc (CD), Digital Video Disc (DVD), or USB storage device, compatible with a standard personal computer. |  |  |  |  |
|  | Patient Results shall be printed out through built-in printer and have an option to print through a connected network printer or external printer. |  |  |  |  |
|  | The analyzer shall have capability for direct interface to connect and transmit data via ASTM or HL7 to hospital LIS. |  |  |  |  |
|  | The analyzer shall be equipped with a barcode reading device capable of reading 1D or 2D barcode system of Code 39 and Code 128. |  |  |  |  |
|  | A built in LED light at the sampling area for operator to easily facilitate sampling under a dim light environment. |  |  |  |  |
|  | The analyzer shall be able to setup operator accounts and define different access right for at least different categories of users: |  |  |  |  |
|  | (a) Supervisor |  |  |  |  |
|  | (b) Key Operator |  |  |  |  |
|  | (c) User |  |  |  |  |
|  | (d) Service |  |  |  |  |
|  | (e) Training |  |  |  |  |
|  | The analyzer shall allow to define different access right for different secure functions including: |  |  |  |  |
|  | (a) Data Management system |  |  |  |  |
|  | (b) Manage Configuration |  |  |  |  |
|  | (c) Lock Analyzer |  |  |  |  |
|  | (d) Lock Analyte |  |  |  |  |
|  | (e) Perform backup and restore |  |  |  |  |
|  | (f) Shutdown and Restart Analyzer |  |  |  |  |
|  | (g) Analyze Sample |  |  |  |  |
|  | (h) View Results |  |  |  |  |
|  | (i) Amend Patient temperature |  |  |  |  |
|  | (j) Accept or Exclude Tests |  |  |  |  |
|  | Quality Management System |  |  |  |  |
|  | The analyzer shall be equipped with an automatic quality management system (AQMS) for active, continuous, and real-time quality assessment process program. |  |  |  |  |
|  | The AQMS shall be capable to detect errors before, during and after sample analysis for each sample. |  |  |  |  |
|  | The AQMS shall provide quality data which shall be collected at the time just before the sample run and just after the sample run, i.e. there shall be at least 2 quality data points for a sample run and there shall be at least 11 data points for 10 sample runs and so on. |  |  |  |  |
|  | The AQMS shall replace every 8-hour requirement of running the traditional external quality controls. |  |  |  |  |
|  | The AQMS shall run more frequent than traditional quality control, and at least 2 levels of QC run for fewer than 8 hours. |  |  |  |  |
|  | The Supplier shall provide a supportive document to demonstrate how the AQMS functions and fulfills the requirements of hospital accreditation. |  |  |  |  |
|  | The AQMS shall have proof for the average error detection time as low as 20 minutes. |  |  |  |  |
|  | The analyzer shall provide immediate error detection and automatic correction. The corrective actions shall be automatically initiated without the need of any manual intervention and documentation. |  |  |  |  |
|  | (a) Automatic corrective actions shall include a dedicated rinse cycle if micro-clots and interferences are detected, no manual intervention is required. |  |  |  |  |
|  | (b) Corrective action documentation shall be fully automatic including confirming errors have been resolved prior to allowing sample testing. |  |  |  |  |
|  | The sample result shall be suppressed on analyzer for reporting automatically when there are reading variations higher than the internal limits during the sample passes over the sensors at measurement. |  |  |  |  |
|  | The analyzer shall provide automatic and continuous checks for the following categories: |  |  |  |  |
|  | (a) System Checks shall routinely conduct functional checks of system components including mechanical, electronics and fluidics before each sample analysis and at various scheduled times throughout the cartridge use-life. |  |  |  |  |
|  | (b) Sensor Checks shall automatically and continuously verify sensor, CO-Ox and cartridge performance throughout the cartridge use-life; |  |  |  |  |
|  | (c) Pattern Checks shall identify the errors caused by microclots and interferences by software pattern/algorithm verification; |  |  |  |  |
|  | - Upon error identification, automatic corrective actions shall be taken and alert the operator. |  |  |  |  |
|  | - The software shall be able to alert the operator with flagging specific to microclots and interferences . |  |  |  |  |
|  | - The data collected for pattern check shall be analyzed based on the voltage output of the measurement sensors. |  |  |  |  |
|  | (d) System shall have solution stability check to verify the reagent stability during the cartridge use-life. |  |  |  |  |
|  | (e) Real time continuous measurement checks during sample analysis shall evaluate sensor performance by abnormal sensor response pattern due to microclots, microbubbles and interferences through slope shape and coefficient values. |  |  |  |  |
|  | - Real time continuous measurement checks shall provide sample integrity checks throughout the entire measurement process starting from sample aspiration to results by taking at least 15 consecutive readings to ensure result accuracy. |  |  |  |  |
|  | - Real time measurement checks shall be able to detect transient sample errors including microclots and interferences |  |  |  |  |
|  | The analyzer shall automatically run quality control checks immediately after the analysis of every sample and as frequent as every 30 seconds. |  |  |  |  |
|  | The analyzer shall have a single cartridge embedded with at least five process solutions for automatically quality checks to verify the cartridge performance. |  |  |  |  |
|  | (a) The analyzer shall disable individual parameter if automatic quality check is failed. |  |  |  |  |
|  | (b) The analyzer shall disable individual parameter if automatic fixing of the particular parameter is failed. |  |  |  |  |
|  | (c) The analyzer shall be able to automatically document all corrective action. |  |  |  |  |
|  | The software shall record all data from the quality check from the five levels of process control solutions. |  |  |  |  |
|  | (a) The analyzer shall perform 2-levels QC check out of the five levels every 4 hours or less without user intervention |  |  |  |  |
|  | (b) The analyzer shall perform 2-levels QC check out of the five levels every 12 hours or less without user intervention. |  |  |  |  |
|  | (c) The software shall display the data in a delta chart format. |  |  |  |  |
|  | (d) The software shall display the data in a monthly format. |  |  |  |  |
|  | (e) The software shall display the data with asterisk or equivalent to indicate corrective action has been done. |  |  |  |  |
|  | (f) The quality control data shall be saved, retrieved, and sent to designated data management software. |  |  |  |  |
|  | (g) The quality control data shall not be able to be edited and deleted for data security. |  |  |  |  |
|  | The corrective actions shall be saved, retrieved and sent to designated data management software. |  |  |  |  |
|  | The event log including the run sample, the aborted test, the removed cartridge and the inserted cartridge shall be saved, retrieved and sent to designated data management software. |  |  |  |  |
|  | Consumables and Accessories |  |  |  |  |
|  | The analyzer shall adopt an all-in-one single closed cartridge containing all the sensors, all process solutions, sample probe and waste bag sharing the same on-board use-life. |  |  |  |  |
|  | (a) The cartridge shall be all-in-one single closed system for ultimate simplicity in maintenance and inventory management |  |  |  |  |
|  | (b) The analyzer shall be an enclosed system with all-in-one cartridge without exposing the operator to biohazardous waste or tubings during maintenance for a good infectious control practice |  |  |  |  |
|  | (c) The sample probe shall not be changed externally and shall not expose the operator to biohazard risk. |  |  |  |  |
|  | All the consumables including all sensors, all process solutions, sample probe and waste bag for a complete sample measurement shall have at least 31 day onboard stability |  |  |  |  |
|  | (a) The sample probe shall be replaced together at the same time with cartridge replacement without additional cost. |  |  |  |  |
|  | (b) The loading of cartridge from one analyser to another shall be blocked if the cartridge has been rejected to guarantee the cartridge performance and traceability. |  |  |  |  |
|  | All involved consumables for complete sample measurement shall have the test size covered from 75 to 600 tests. |  |  |  |  |
|  | (a) The test size shall be true reportable test count which QC and Calibration shall not be included. |  |  |  |  |
|  | The total warm-up time including both off-board and on-board of the all involved consumables for complete sample measurement shall be no more than 40 minutes. |  |  |  |  |
|  | All consumables for complete sample measurement including all sensors, all process solutions, sample probe and waste bag, shall be able to be stored at room temperature (15oC -25oC). |  |  |  |  |
|  | Quality control and Calibration is applied automatically and no sample test number is deducted, ensuring the accuracy of the results with no extra consumption on sample test numbers. |  |  |  |  |
|  | The analyzer including measuring sensors shall be free of maintenance, no filling and condition solution, and no change of electrode is required. |  |  |  |  |
|  | The analyer shall not require adaptor for syringe, capillary, tube, and ampoule samples. |  |  |  |  |
|  | The analyzer shall allow user to manually adjust the depth of the probe into the sample inside the syringe in order to reduce the sample dead volume. |  |  |  |  |
|  | Minimum sample requirements for the cartridge shall be less than or equal to 150uL. |  |  |  |  |
|  | Time to results shall be as quick as 45 seconds from sample introduction. |  |  |  |  |
|  | The components and processes used to manufacture the process control solutions in consumable pack are traceable to NIST standards. |  |  |  |  |
|  | User Interface, Connectivity, Data Security |  |  |  |  |
|  | The analyzer shall have the capability of direct interface or connect through middleware to LIS/HIS. |  |  |  |  |
|  | The analyzer shall have at least 4 serial communication ports and at least 1 TCP/IP port for connectivity purpose. |  |  |  |  |
|  | The analyzer shall allow the data transmission through the serial communication port and the TCP/IP port simultaneously. |  |  |  |  |
|  | The analyzer shall have capability to store test results uploaded from analyzers in case the LIS is down or during maintenance and the stored test results shall be uploaded automatically when LIS connectivity is resumed |  |  |  |  |
|  | Supplier shall be responsible for the interfacing of analyzer to the LIS and ensure the bi-directional transmission of order and result. |  |  |  |  |
|  | The QC data could be download to the computer or third party software to review of QC and assess the performance of analyzer. Supplier shall provide technical support for the transmission of data. |  |  |  |  |
|  | Regardless of where the problems may arise, the Successful Supplier must provide all the necessary assistance, consultancy, time and resources, including the replacement of hardware, as requested by the hospital to ensure the successful completion of the LIS/EQUIPMENT interface |  |  |  |  |
|  | Electricity and Safety Requirement |  |  |  |  |
|  | The Supplier shall provide uninterrupted power supply (UPS) of sufficient capacity to support the system for not less than 15 minutes during power failure at no extra charge. |  |  |  |  |
|  | Supplier shall provide startup reagents and consumables for the evaluation of analyzer performance in accordance with ISO:15189 requirement. It is also one of the criteria of acceptance of the analysers. |  |  |  |  |
| B. | Implementation Services |  |  |  |  |
|  | The System shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included within 10 weeks from the date informed by the Hospital. |  |  |  |  |
|  | The Supplier shall be responsible for connecting all water inlets, waste outlets and electricity supply to the System and any modification works if required. |  |  |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |  |  |
|  | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |  |  |
|  | The Supplier shall be responsible for the make good of ceiling, wall and floor in the installation site after completion of equipment installation. |  |  |  |  |
| C | Training |  |  |  |  |
|  | On-site maintenance and operational training shall be provided at no additional charges for a minimum of twice of each in the first two years when required. |  |  |  |  |
|  | The supplier shall be responsible to provide session of on-site maintenance training to representatives of CMHHK upon request. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures , calibration and alignment, adjustment. |  |  |  |  |
|  | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. Detailed syllabuses shall be submitted for approval, upon request. The practical part of the training shall coincide with the installation and commissioning of the Goods. |  |  |  |  |
|  | Training protocol of operator and content shall be submitted together with the tender return for evaluation Multiple operator training shall be required upon the request of laboratory staff and certificate shall be provided after the training as the reqirement of laboratory accreditation. |  |  |  |  |
|  | Additional on-site application sessions shall also be arranged upon request by users during the warranty period without additional cost. |  |  |  |  |
| D | Documentation |  |  |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |  |  |
|  | At the time of delivery of the equipment, appropriate Set(s) of the manufacturer’s original operation and maintenance manuals in English or in Chinese complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to the hospital after order placement but before the equipment delivery for documentation purpose. |  |  |  |  |
| 4. | The Supplier shall provide the authorization letter from manufacturer for providing products, parts and service to end user. The written undertaking shall be signed by a duly authorized representative of the manufacturer and dated no later than the quotation closing date. |  |  |  |  |
| E | Acceptance Tests |  |  |  |  |
|  | Safety Test For the purpose of this contract the Goods shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Authority Representative or his authorized nominee or nominees. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the Goods. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |  |  |
|  | Functional Test For the purpose of this Contract the Goods shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. The successful tenderer shall provide all reagents and consumables for the evaluation. In the event that the equipment fails to conform to the above stated requirements, the successful tenderer is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the Goods shall be determined by the Hospital based upon the satisfactory completion of such functional test. |  |  |  |  |
|  | On-site Functional Test The Supplier shall provide on-site free technical support by the qualified product specialist(s) and all the required reagents and any necessary consumables for functional test, evaluation, initial startup, correlation and customization of the system without any additional cost. |  |  |  |  |
| F | Indicative Warranty Service |  |  |  |  |
|  | The potential supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The potential supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |  |  |
|  | The potential supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least two times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |  |  |
|  | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |  |  |
|  | The potential supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure which may arise within the warranty period of the equipment. |  |  |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the supplier to carry on working until the particular maintenance task is completed. |  |  |  |  |
|  | Upon completion of the corrective maintenance works, the potential supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests (Please refer to section E in Part 3 for details) |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Whole Blood Analyzers can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Whole Blood Analyzers does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Whole Blood Analyzers in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company.  Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 2 units |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

(b) Indicative Price Information for Selected Desirable Features (if applicable)

| Aspect | Description of Selected Desirable Features | Any Additional Charge to Total One-time Charge as Specified in Part 7(a) (Please tick whichever is applicable) |
| --- | --- | --- |
| Examples: |  |  |
| 1 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 1 hour < continuous operation ≤ 2 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 2 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 2 hours < continuous operation ≤ 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 3 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: continuous operation > 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |

Part 8 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END