P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Baths, Dry

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Ms Teresa TONG)

### [by fax: 21274795 or email: tlmtong@healthbureau.gov.hk]

Your ref: -

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and installation of Baths, Dry (hereinafter refers as the “System”) for the Chinese Medicine Hospital of Hong Kong (“opCMHHK”) located at 1 Pak Shing Kok in Tseung Kwan O, New Territories, Hong Kong.  The CMHPO therefore wishes to collect market information on this System.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Baths, Dry that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Baths, Dry.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ____________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “System” in this Part 3 refers to section A1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient)

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |
|  | General Description |  |  |
|  | Two (2) sets of dry baths block shall be supplied, delivered and installed in the Patholoogy Laboratory of CMHHK. |  |  |
|  | Functional and Performance Requirements of the System |  |  |
|  | The dry baths shall feature a compact design suitable for benchtop use in a clinical laboratory. |  |  |
|  | The dry baths shall be made of powered-coated steel. |  |  |
|  | The dry baths shall be compatible with various interchangeable blocks, accommodating different tube sizes and types, including 12x75mm glass or plastic tubes, microcentrifuge tubes, PCR tubes, and other commonly used laboratory tubes. |  |  |
|  | The dry baths shall feature a digital display and controls for real-time monitoring and setting of temperature and time. |  |  |
|  | The control panel shall include digital display with touchpad or knob for setting temperature and time. |  |  |
|  | The control panel shall feature a start/stop button for initiating and pausing the heating process. |  |  |
|  | The dry baths shall include a built-in timer with a range of 1 minute to 99 hours and 59 minutes, or a wider range. |  |  |
|  | The dry baths shall have a heating indicator light to show when the heating process has started. |  |  |
|  | The dry baths shall operate within a temperature range of ambient +5°C to 130°C (± 0.2°C). |  |  |
|  | When operating at 37°C, the dry baths shall offer a temperature accuracy of ± 0.5°C or better. |  |  |
|  | When operating at 37°C, the dry baths shall maintain a temperature uniformity of ±1.0°C or better across the block. |  |  |
|  | The temperature control system shall be precise and based on a Proportional Integral Derivative (PID) circuit which is a control loop feedback mechanism to regulate temperature. |  |  |
|  | The dry baths shall be equipped with an over-temperature protection system to prevent overheating. |  |  |
|  | Accessories |  |  |
|  | Each set of bath shall include interchangeable blocks capable of accommodating: |  |  |
|  | 1.5 ml microtubes for at least 28 tubes |  |  |
|  | 2.0 ml microcentrifuge tubes for at least 28 tubes |  |  |
|  | 13 mm tube for at least 20 tubes |  |  |
|  | The interchangeable blocks shall be made of aluminium or other high thermal conductivity materials providing good heat conductivity. |  |  |
|  | The interchangeable blocks shall be easy to clean by a damp cloth. |  |  |
|  | Weight and Dimensions |  |  |
|  | The weight of the bath shall be less than 5.0 kg. |  |  |
|  | The dimensions of each bath shall be ranging within the following: - Length: 315 mm to 330 mm; - Width: 200 mm to 250 mm; - Height: 100 mm to 125 mm. |  |  |
|  | Safety and Product Standards |  |  |
|  | The dry baths shall comply with the following international standards (or equivalent national and other recognised standards or certifications): |  |  |
|  | IEC 61010-1:2010 Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements; and |  |  |
|  | Certification to meet relevant safety and performance standards (e.g. CE, UL certifications). |  |  |
|  | Electricity Supply and Environmental Conditions |  |  |
|  | The dry baths shall be suitable for operation at 220 volt (± 6%), 50 Hz (± 2%). |  |  |
|  | The dry baths shall be suitable for operation at an ambient temperature of 10°C - 30°C, ≤ 80% relative humidity. |  |  |
|  | Implementation Services |  |  |
|  | The System shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included within 10 weeks from the date informed by CMHHK. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
|  | Training |  |  |
|  | The Supplier shall be responsible to provide session of on-site maintenance training to representatives of CMHHK upon request at no additional charges. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures, calibration and alignment, adjustment. |  |  |
|  | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. |  |  |
|  | Documentation |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |
|  | At the time of delivery of the equipment, appropriate Set(s) of the manufacturer’s original operation and maintenance manuals in English or in Chinese complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to CMHHK after order placement but before the equipment delivery for documentation purpose. |  |  |
|  | The Supplier shall provide the authorization letter from manufacturer for providing products, parts and service to end user. The written undertaking shall be signed by a duly authorized representative of the manufacturer and dated no later than the quotation closing date. |  |  |
|  | Acceptance Tests |  |  |
|  | Safety Test For the purpose of this contract the System shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Contractor with the witness of representative of CMHHK. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the System. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
|  | Functional Test For the purpose of this Contract the System shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. In the event that the equipment fails to conform to the above stated requirements, the successful tenderer is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the Systems shall be determined by CMHHK based upon the satisfactory completion of such functional test. |  |  |
|  | Indicative Warranty Service |  |  |
|  | The Supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The potential Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
|  | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least one times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |
|  | The Supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure which may arise within the warranty period of the equipment. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the potential Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
|  | Indicative Maintenance Service |  |  |
|  | The Supplier shall quote the charge for annual maintenance services after the warranty period within the serviceable life of the proposed System. |  |  |
|  | The Supplier shall submit a price list of all spare parts of the System chargeable to the CMHHK Operator. For spare parts not covered by the submitted prices, the potential Supplier must submit a quotation to the CMHHK Operator for consideration every time when spares are required. |  |  |
|  | The Supplier shall deploy properly trained service personnel to carry out the maintenance services and shall ensure that all necessary precautions for their safety are taken. |  |  |
|  | The Supplier shall provide free of additional charge corrective maintenance service for providing immediate repair service for the System and related equipment in normal working hours. |  |  |
|  | The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the potential Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work is not completed at the end of particular normal working period, subject to the user’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
|  | Spare Parts |  |  |
|  | The Supplier shall guarantee the availability of maintenance spare parts for the anticipated life of the System. Sufficient spare parts shall be held by the successful Supplier to cater for the maintenance during the warranty period. |  |  |
|  | The Suppliers, in their tender submission, shall provide a comprehensive list of recommended spare parts with unit prices valid for at least one (1) year after expiry of warranty. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Baths, Dry can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Baths, Dry does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Baths, Dry in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| ## IEC61010-1:2010 or equivalent | ## Details please refer to the Section 5.1.1 in Part 3 |  |  |  |
| CE certification or equivalent | Details please refer to the Section 5.1.2 in Part 3 |  |  |  |
| UL certification or equivalent | Details please refer to the Section 5.1.2 in Part 3 |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-month warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge* (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

Note: * The Total One-time Charge shall include one-year of warranty period.

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-month period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-month period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END