P R O F O R M A

Request for Market Information (“RFI”) for 
Supply and Installation of Microscope and Laboratory

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Ms Wing LI, EM(CMHPO)1B)

### [by fax: 2127 4795 or email: wlfli@healthbureau.gov.hk]

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and installation of Microscope and Laboratory (hereinafter refers as the “Goods”) for the Chinese Medicine Hospital of Hong Kong (“CMHHK”) located at 1 Pak Shing Kok in Tseung Kwan O, New Territories, Hong Kong.  The CMHPO therefore wishes to collect market information on this Goods.

- Background of the CMHHK Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Microscope and Laboratory that may meet the requirements of the Goods stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Microscope and Laboratory.

Part 2 – General Information of the Goods

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the Goods |  |
| - Model number/ name/ version number of the Goods |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the Goods (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the Goods (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life) | The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components: ____________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “Goods” in this Part 3 refers to section A below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed Goods “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed Goods in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed baths, dry against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |
| 1 | Goods to be Supplied |  |  |
| 1.1 | It calls for the supply and installation of four units of Microscope for laboratory use. (hereinafter refers as “Microscopes” or “Units”). |  |  |
| 1.2 | (a) two units of Upright general microscope |  |  |
| 1.3 | (b) two units of identical Upright microscope (a) with additional cameras and photographic function |  |  |
| 2 | Intended Use of the Goods |  |  |
| 2.1 | The Laboratory microscopy is versatile tool in various laboratory procedures and commonly used for blood smear examination, morphologyical studies of cells. |  |  |
| 2.2 | With the carmera and photographic function, it enables the sharing of image for education, discussion and consultation purposes. |  |  |
| 3 | General Description |  |  |
| 3.1 | A microscope with tilting / telescoping eyepiece tube and with optical system to accommodate various accessories to meet different application. |  |  |
| 3.2 | The microscope shall have ergonomic design to ensure ideal viewing posture, enabling long hours of observations to be carried out in comfort. |  |  |
| 3.3 | The microscope shall provide optical system so that multiple optical accessories can be added in between ocular tube and objectives. |  |  |
| 3.4 | All components along the microscope should be come from the same brand and manufacturer to make sure the microscope achieved optimium performance, compatibility and stablility. |  |  |
| 4 | Microscope Main Body |  |  |
| 4.1 | The base of stand shall be sturdy and stable with rigidity to accommodate accessories for upgradeability. |  |  |
| 4.2 | The field diaphragm, power switch and illumination system shall be ergonomically placed at low position within easy reach. |  |  |
| 4.3 | The nosepiece with at least 5 objective positions. |  |  |
| 4.4 | At least five objectives, including 4X, 10X, 20X, 40X, and 100X (oil immersion), shall be fixed to main body. |  |  |
| 4.5 | Built-in digital display to show easily the magnification and brightness setting status |  |  |
| 4.6 | Coarse movement of control knob shall not less than 13mm per rotation and the fine movement shall not more than 0.1mm per rotation. |  |  |
| 4.7 | Fine focusing knob shall be graduated in one micron or finer increment. |  |  |
| 4.8 | Torque adjustment shall be available for coarse focusing knob. |  |  |
| 4.9 | Stage upper limit clamp shall be provided to eliminate the crash of slides or damage of the objectives. |  |  |
| 4.10 | The position of the camera port shall be designed at rear side of ocular tube so that no additional intermediate tube is required. |  |  |
| 4.11 | The microscope stage is adjustable up and down for object focus in not less than 25mm range. |  |  |
| 5 | Illumination System |  |  |
| 5.1 | The illuminator shall use a White LED, or an equivalent light source. |  |  |
|  | The white LED with high color rendering index has wavelength characteristics close to that of halogen, realizing halogen lamp-equivalent color reproductions. LED lighting retains this color temperature to prevent images from turning reddish or bluish in color. |  |  |
| 5.2 | There is built-in lens array design which provide best and even images from low power to high power objectives. |  |  |
| 5.3 | There is built-in light intensity management function automatically stores brightness settings between different objective lens during observation. It eliminates the need to manually adjust the illuminator brightness every time switched objective lenses. |  |  |
| 5.4 | The microscopy shall automatically turn off the illumination after a certain period of inactivity, saving electricity and preventing the illumination from being left on. |  |  |
| 5.5 | A light intensity switch shall be provided and separate from the power switch. |  |  |
| 6 | Mechanical Stage |  |  |
| 6.1 | The mechanical stage shall incorporate a super-hard alumite coated surface to increase durability & smoothness. |  |  |
| 6.2 | The microscopes shall provide a rectangle surface not less than 150mm x 200mm. |  |  |
| 6.3 | The microscopes shall provide cross travel (X-Y movement) of not less than 70mm by 50mm. |  |  |
| 6.4 | It shall be low in position with the height of stage 200mm ± 10mm from the base for ergonomic operation. |  |  |
| 6.5 | It shall be controlled by coaxial X & Y control knob on the right hand side. |  |  |
| 6.6 | The stage shall accommodate 2 glass slides (12x75mm) at a time. |  |  |
| 6.7 | Motion torque for X & Y control of the stage shall be adjustable. |  |  |
| 6.8 | Stage handle height and torque adjustable without tools shall be provided for enhancement comfort during operation. |  |  |
| 7 | Eyepiece Tube and Eyepieces |  |  |
| 7.1 | lnterpupillary distance of eyepieces shall be adjustable. |  |  |
| 7.2 | Eyepiece tube shall provide an inclination angle from 10° to at least 25° upward from horizontal. |  |  |
| 7.3 | Eyepiece tube shall provide a built-in sliding mechanism of not less than 40mm movement. |  |  |
| 7.4 | The eye-level shall be adjustable by the riser for ergonomic features. |  |  |
| 7.5 | Eyepiece tube shall incorporate two 10X eyepieces with diopter adjustment. |  |  |
| 7.6 | Eyepieces shall be focusable with height of eye-point 330mm ± 10mm and chromatic aberration free. |  |  |
| 7.7 | The eyepieces shall provide field of view of 22mm or more |  |  |
| 7.8 | Built-in camera port at the back of tilting eyepiece tube. Digital camera can be adapted without rise up the eye Level |  |  |
| 8 | Condenser Lens & Optics |  |  |
| 8.1 | The condenser lens shall be of swing out type with Numerical Aperture (N.A.) not less than 0.90 |  |  |
| 8.2 | The condenser shall be designed for use with 1X to 100X objectives or better. |  |  |
| 8.3 | Interchanging objectives shall not disturb the parfocality. |  |  |
| 8.4 | Objective lens shall be of Plan Semi Apochromat or compatible Chromatic Aberration free Infinity types. |  |  |
| 8.5 | One 4X objective of N.A. at least 0.10 and working distance (W.D.) 17.0mm or better shall be provided. |  |  |
| 8.6 | One 10X of N.A. at least 0.25 and W.D. 10.0mm or better shall be provided. |  |  |
| 8.7 | One 20X objective of N.A. at least 0.5 and W.D. 1.2mm or better shall be provided. |  |  |
| 8.8 | One 40X objective of N.A. at least 0.75 and W.D. 0.56 mm or better shall be provided. |  |  |
| 8.9 | One 100X, oil immersion, objective of N.A. 1.3 and W.D. 0.16mm or better shall be provided. |  |  |
| 9 | Digital camera, for section A1.3(b) only |  |  |
| 9.1 | The camera head shoud be installed microscope body without rise up the eye Level |  |  |
| 9.2 | The camera system should work alone without Personal Computer (PC) connection. |  |  |
| 9.3 | The camera system should support HDMI2.0 interface connection to external displays for synchronous output, 4K/1080P auto switching according to monitor resolution. |  |  |
| 9.4 | The camera system should provide real-time video Eextend Depth of Focus function. |  |  |
| 9.5 | The camera system should provide real-time image stitching to capture larger sample area. |  |  |
| 9.6 | The camera system should have built-in 5G WiFi connection function for camera control or image capture by smart devices |  |  |
| 9.7 | The camera system should included free apps for Android and iOS devices |  |  |
| 9.8 | The camera should come with free PC software to control the camera and image processing, supporting measurement functions. |  |  |
| 9.9 | The Sensor Type is Sony IMX412 CMOS sensor or better |  |  |
| 9.10 | The Image sensor shall be >= 1/2.3 inch Color CMOS image sensor. |  |  |
| 9.11 | The Recordable pixels shall be at least 8 megapixels, 3840 x 2160 pixels |  |  |
| 9.12 | The Live display speed (Maximum fps) shall be at least 3840 x 2160 (8 megapixels) or faster and match with recordable/ display pixels. |  |  |
| 10 | Safety and Product Standards |  |  |
| 10.1 | The Goods shall comply with the following international standards (or equivalent national and other recognised standards or certifications): |  |  |
| 10.2 | IEC 61010-1:2010 Safety requirements for electrical equipment for measurement, control, and laboratory use - Part 1: General requirements; and |  |  |
| 10.3 | Certification to meet relevant safety and performance standards (e.g. CE, UL certifications). |  |  |
| 11 | Electricity Supply and Environmental Conditions |  |  |
| 11.1 | The Goods shall be suitable for operation at 220 volt (± 6%), 50 Hz (± 2%). |  |  |
| 11.2 | The Goods shall be suitable for operation at an ambient temperature of 10°C - 30°C, ≤ 80% relative humidity. |  |  |
| 12 | Other Requirement |  |  |
| 12.1 | The unit shall contain all acceessories such as one vinyl cover, one power cord for 220V. |  |  |
| 12.2 | The price quoted shall include local delivery and installation within 12 weeks upon receipt of confirmed order. |  |  |
| B | Implementation Services |  |  |
| 1 | The Goods shall be tested and become ready for use by the timeline specified in Part 4(k) with all costs included. |  |  |
| 2 | The successful supplier shall be responsible for connecting all the electricity supply to the Goods and any modification works if required. |  |  |
| 3 | The successful supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery of the Goods at no charges. |  |  |
| 4 | The successful supplier shall note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during Goods transportation. The successful supplier shall take all due measures to protect such property. |  |  |
| 5 | The Goods delivery and acceptance tests must be completed within 4 weeks from the date informed by the hospital. |  |  |
| C | Training |  |  |
| 1 | The successful supplier shall provide basic operation training and maintenance training, with the necessary training material, within the first year operation of the Goods at no additional cost. |  |  |
| 2 | Basic operation and maintenance training shall be at least twice per year in the first year of operation of the Goods. |  |  |
| 3 | The training shall include, but not limited to, the theory and technique in the operation and setting up of the Goods, on-site practical of various test profiles, troubleshooting, fault diagnosis, preventive maintenance procedures, use of special tools and test instruments. |  |  |
| 4 | Additional on-site application sessions shall also be arranged upon request by users during the warranty period without additional cost. |  |  |
| D | Documentation |  |  |
| 1 | Two complete sets of the manufacturer’s operation and maintenance manuals shall be submitted for each Goods at the time of delivery of the Goods. The manufacturer’s operation manual shall be submitted in one hardcopy and one softcopy. The manufacturer’s maintenance manual shall be provided in two softcopies. |  |  |
| 2 | The maintenance manuals and checklist for preventive maintenance of the Goods shall be sent to the hospital after order placement but before the Goods delivery for documentation purpose. |  |  |
| E | Acceptance Tests |  |  |
| 1 | The complete set of Goods shall be subject to an acceptance test by the hospital department after delivery has been completed by the successful supplier. |  |  |
| 2 | The hospital will not conduct acceptance tests until it has received full certification, as specified in Section A of this part, from the successful supplier. The certificates must confirm that the successful supplier has conducted its own tests and found that the Goods meets the specifications in the Contract. |  |  |
| 3 | Should the Goods fail to comply with the safety standards listed in above clause, the successful supplier shall investigate the causes of the failure and take appropriate corrective actions. The hospitals must be immediately informed of the failure. |  |  |
| F | Indicative Warranty and Maintenance Service |  |  |
| 1 | The successful supplier shall provide at least 12-months warranty period, starting from the date of acceptance of the system. During this period, all services which include replacement of faulty parts, scheduled and breakdown services by qualified maintenance personnel, shall be free of charge. |  |  |
| 2 | The successful supplier shall be responsible for the provision of on- going preventive maintenance (PM) and corrective maintenance (CM) services throughout the warranty period by trained service engineer. |  |  |
| 3 | During the warranty period, PM services shall be performed at least twice per year and the Goods’ correlation reports that meet the requirements of the local accreditation body shall be provided after each PM. |  |  |
| 4 | Service shall be performed by manufacturer-trained service engineers and stay locally for the on-call services. |  |  |
| 5 | During the warranty period, CM service call shall be unlimited and shall be carried out within 4 hours for emergency calls and no longer than 8 hours for other calls. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The Goods should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the Goods (if applicable) |  |  |
|  | #### Delivery of the Goods |  |  |
|  | #### Installation of the Goods |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### Goods Ready for Use (i.e. the date when the Goods has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed bath, drys can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed bath, drys does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed bath, drys in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing

(Not applicable)

(Note to Suppliers: Please advise whether your company and the proposed Goods have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the Goods | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed Goods been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed Goods have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed Goods has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed Goods have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed Goods has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed Goods have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed Goods been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed Goods (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the Goods

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the Goods and related accessories, as more particularly specified in section A in Part 3, including the provision of a minimum 12-months warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the Goods, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge* (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

Note: * The Total One-time Charge shall include one-year of warranty period.

(b) Indicative Price Information for Selected Desirable Features (Not applicable)

| Aspect | Description of Selected Desirable Features | Any Additional Charge to Total One-time Charge as Specified in Part 7(a) (Please tick whichever is applicable) |
| --- | --- | --- |
| Examples: |  |  |
| 1 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 1 hour < continuous operation ≤ 2 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 2 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 2 hours < continuous operation ≤ 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 3 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: continuous operation > 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed Goods shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the Goods may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed Goods is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service.

- Indicative Maintenance Prices of the Proposed Goods

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of Goods’ Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the Goods’ software during the serviceable life of the Goods for the CMHHK Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the Goods; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed Goods Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed Goods (leave blank if information is not available)

My/our proposed Goods was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed Goods (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the Goods, such as the requirements of ceiling mount support, power supply requirements, etc.)

END