P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Plasma-thawing Device

to the Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Mr S.L. CHENG)

### [by fax: 2127 4795 or email: slcheng@healthbureau.gov.hk]

Your ref:  --

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 8 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply of Plasma-thawing Device (hereinafter refers as the “Goods”) for the Chinese Medicine Hospital of Hong Kong (“opCMHHK”) located at 1 Pak Shing Kok in Tseung Kwan O, New Territories, Hong Kong.  The CMHPO therefore wishes to collect market information on this systems.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Plasma-thawing Device that may meet the requirements of the Goods stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Plasma- thawing Device.

Part 2 – General Information of the Goods

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the Goods |  |
| - Model number/ name/ version number of the Goods |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the Goods (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the Goods (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life) | The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components: ____________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “Goods” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed Goods “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed Goods in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |
|  | General Description |  |  |
|  | Two (2) sets of plasma thawing device shall be provided with design for use in hospital blood bank and clinical laboratories offering a safe thawing conditions for donor plasma units without direct contact with water for preventing contamination through the process. |  |  |
|  | The Goods shall be in a thawing capacity of accommodating one (1) to four (4) units of donor frozen plasma bags simultaneously. |  |  |
|  | The Goods shall provide an average thawing time of less than 20 minutes at 370C for four (4) units of 450 mL plasma products. |  |  |
|  | The Goods shall comply with the international standards for medical devices (e.g. ISO13485, FDA approval, CE marking or equivalent) and meet the needs of the blood transfusion guidelines in Hospital Authority locally. |  |  |
|  | Construction |  |  |
|  | The Goods shall be equipped with a warming chamber with a cover. |  |  |
|  | The cover of the warming chamber shall meet the following requirements:- - the cover shall be closed to retain heat during thawing process; |  |  |
|  | (b) the cover shall be transparent for easy visual inspection of the donor units and heating cushions inside the chamber; |  |  |
|  | (c) the cover shall be equipped with a lock/release device for safety issues. |  |  |
|  | The warming chamber shall contain at least two (2) warming cushions and a paddle for warming donor blood products. |  |  |
|  | The two warming cushions shall meet the following requirements: (a) the cushions shall be soft as better contact with blood units for heat transfer; |  |  |
|  | (b) the cushions shall be impermeable and resistant for fluid peneatration. |  |  |
|  | The Goods shall be equipped with a paddle for agitating the blood products and fluids to homogenize and facilitate fast thawing. |  |  |
|  | The Goods shall be equipped with a leakage sensor in the warming chamber for the detection of any fluid leakage, e.g. donor blood compponents or water. |  |  |
|  | The warming chamber shall have a shallow design for easy viewing and cleaning when leakage occurs. |  |  |
|  | Digital or analog temperature controls device with LCD or LED display for real-time monitoring of temperature and thawing progress. |  |  |
|  | The operating panel shall have large display with plain text. |  |  |
|  | The Goods shall have an opening for filling up the device with heat transfer fluid. It shall have a screw plug to cover the opening. |  |  |
|  | The Goods shall have a tap with tubing at the bottom of the device for drainage of the heat transfer fluid inside the device. |  |  |
|  | The Goods shall have smooth surfaces and round edges for safety and easy cleansing. |  |  |
|  | The Goods shall be made of corrosion resistant materials. |  |  |
|  | The Goods shall be equipped with a temperature sensors for periodic calibration checking. |  |  |
|  | Operation of the Goods |  |  |
|  | The blood component bags shall be loaded horizontally in the warming chamber. It allows easy visual inspection of the blood product status and any leakage. |  |  |
|  | The warming cushions shall be securely locked in the warming chamber, and easily detachable with a special device or tools with the system. |  |  |
|  | When filling the Goods with heat transfer fluid, electronic sensor shall detect the fill level, and hints shall be shown on the display monitor. |  |  |
|  | The heat transfer fluid shall flow in the heater and the warming cushions. |  |  |
|  | The Goods shall have at least 4 warming program presets and 1 user configurable warming program. |  |  |
|  | At least two warming program shortcuts can be set onto the home screen by user according to preference. |  |  |
|  | When a warming program is started, the display shall show the program name, the current heat transfer fluid temperature, and the remaining time, among other information. |  |  |
|  | When a warming program is started, the heat transfer fluid shall be heated up, and the warming cushions shall be filled with the heat transfer fluid. |  |  |
|  | While a warming program is running, the user can alter the remaining time - usually between 30 seconds and 99 minutes with selection button ("- or +" or equivalent). |  |  |
|  | If the warming chamber cover is opened while a warming program is in progress, the remaining time shall be 'frozen'. A message shall appear on the display and heat transfer fluid circulation shall be paused. |  |  |
|  | When the warming program is completed, the Goods shall have visual and acoustic signal to alert users. |  |  |
|  | The warming program can be stopped anytime by pressing the “Stop” button. |  |  |
|  | The warming temperature shall be adjustable in every 0.5°C, from +37°C to +45°C. |  |  |
|  | The Goods shall have dual over-temperature protection, by electronic controlled and / or software controlled. |  |  |
|  | When an over-temperature limit is reached, the heating of the Goods shall be switched off, the LED and operating panel shall light up, and a continuous alarm tone shall sound to alert users. |  |  |
|  | When leakage is detected by the moisture sensors, an intermittent acoustic signal shall be triggered, and a warning message shall flash in the display monitor. |  |  |
|  | When the Goods error occurs, a message shall be shown on the display monitor, the fault lamp (LED) on the operating panel shall be activated, and a continuous warning tone shall be triggered. |  |  |
|  | All moving parts shall be safely enclosed inside the warming chamber with cover, for protecting the user and the parts of the Goods. |  |  |
|  | Technical Data |  |  |
|  | The Dimensions of the Goods with the cover closed shall not exceed W350 mm x D600 mm x H350 mm. |  |  |
|  | The Dimensions of the Goods with the cover open shall not exceed W350 mm x D600 mm x H750 mm |  |  |
|  | The capacity of loading of blood product shall be either 4 bags of each ≤ 500 mL or 2 bags of each ≤ 1000 mL. |  |  |
|  | The filling quantity of heat transfer fluid in the Goods shall not exceed 10 L. |  |  |
|  | The weight of the Goods without filling of heat transfer fluid shall not exceed 20 kg. |  |  |
|  | The Goods shall allow a selectable temperature setting from +37°C to +45 °C. |  |  |
|  | The accuracy of the set temperature point shall be ± 0.5 °C. |  |  |
|  | The Goods shall provide an automatically over-temperature protection to alert the user or stop the overheating device if the temperature exceeds pre-set safe limits. |  |  |
|  | Electrical & Safety Requirements |  |  |
|  | The Goods shall remain operational and within specifications throughout the voltage range of 220V ± 6%, 50 HZ ±2%, 1-phase A.C. electrical supply. |  |  |
|  | Single phase mains-operated equipment shall be fitted with a power plug suitable for the site condition. The plug shall comply with relevant standards e.g. BS1363 for 13A plug. |  |  |
|  | The Goods shall comply with the safety requirements of IEC60601-1 / IEC61010-1 or equivalent. |  |  |
|  | The Goods shall comply with the electromagnetic compatibility (EMC) requirements of IEC60601-1-2 / IEC61326-1 or equivalent. |  |  |
| B | Implementation Services |  |  |
|  | The Goods shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included within 10 weeks from the date informed by the Hospital. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
| C | Training |  |  |
|  | The Supplier shall be responsible to provide session of on-site maintenance training to representatives of CMHHK upon request at no additional charges. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures, calibration and alignment, adjustment. |  |  |
|  | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. |  |  |
| D | Documentation |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |
|  | At the time of delivery of the Goods, appropriate Set(s) of the manufacturer’s original operation and maintenance manuals in English or in Chinese complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to the CMHHK after order placement but before the equipment delivery for documentation purpose. |  |  |
| 4. | The Supplier shall provide the authorization letter from manufacturer for providing products, parts and service to end user. The written undertaking shall be signed by a duly authorized representative of the manufacturer and dated no later than the quotation closing date. |  |  |
| E | Acceptance Tests |  |  |
| 1. | Safety Test For the purpose of this Contract the Goods shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Supplier with the witness of representative of CMHHK.. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the Goods. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
| 2. | Functional Test For the purpose of this Contract the Goods shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. The Supplier shall provide all consumables that required for the evaluation, if necessary. In the event that the equipment fails to conform to the above stated requirements, the Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire Goods, where deemed necessary. The date of acceptance of the Goods shall be determined by the CMHHK based upon the satisfactory completion of such functional test. |  |  |
| F | Indicative Warranty Service |  |  |
| 1. | The Supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The potential Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
| 2. | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least one times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
| 3. | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |
| 4. | The Supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure which may arise within the warranty period of the equipment. |  |  |
| 5. | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the supplier to carry on working until the particular maintenance task is completed. |  |  |
| 6. | Upon completion of the corrective maintenance works, the potential Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The Goods should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the Goods (if applicable) |  |  |
|  | #### Delivery of the Goods |  |  |
|  | #### Installation of the Goods |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### Goods Ready for Use (i.e. the date when the Goods has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Plasma-thawing Device can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Plasma-thawing Device does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Plasma-thawing Device in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed Goods have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the Goods | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed Goods been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed Goods have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed Goods has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed Goods have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed Goods has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed Goods have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed Goods been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed Goods (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the Goods

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the Goods and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the Goods, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge* (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

Note: * The Total One-time Charge shall include one-year of warranty period.

(b) Indicative Price Information for Selected Desirable Features (if applicable)

| Aspect | Description of Selected Desirable Features | Any Additional Charge to Total One-time Charge as Specified in Part 7(a) (Please tick whichever is applicable) |
| --- | --- | --- |
| Examples: |  |  |
| 1 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 1 hour < continuous operation ≤ 2 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 2 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 2 hours < continuous operation ≤ 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 3 | The Goods shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: continuous operation > 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |

Part 8 – Supplementary Information

- Number of proposed Goods Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed Goods (leave blank if information is not available)

My/our proposed Goods was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed Goods (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the Goods, such as the requirements of ceiling mount support, power supply requirements, etc.)

END