P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Centrifuge

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

To : 	Project Director (CMHPO)

(Attn. Ms Teresa TONG)

### [by fax: 21274795 or email: tlmtong@healthbureau.gov.hk]

Your ref: -

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and installation of Centrifuge (hereinafter refers as the “System” or “Centrifuge”) for the Chinese Medicine Hospital of Hong Kong (“opCMHHK”) located at 1 Pak Shing Kok in Tseung Kwan O, New Territories, Hong Kong.  The CMHPO therefore wishes to collect market information on this System.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK.  The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong.  It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services.  The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021, and commenced hospital services by phases since 11 December 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Centrifuge that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Centrifuge.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ____________________________________ (Please also provide the expected life of these excluded components) |
| - *Total weight of the proposed System | ____________kg |

*The maximum floor loading capacity where the System is to be installed is 5 kPa. Please ensure that your proposed System can comply with this requirement.

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “System” in this Part 3 refers to section A1.1-1.7 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient)

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |  |
|  | Overall Requirements |  |  |  |
| 1.1. | System to be Supplied |  |  |
| 1.1.1 | The document calls for the supply and installation of multi-purpose centrifuge (hereinafter refers as “System” or “Centrifuge”). |  |  |
| 1.2. | The System shall have the following items:- |  |  |
| 1.2.1. | Two (2) sets of benchtop, high-speed centrifuge as detailed in section A2 below; |  |  |
| 1.2.2. | Two (2) sets of benchtop, low-speed, non-refrigerated centrifuge as detailed in section A3 below; |  |  |
| 1.2.3. | One (1) set of benchtop, low-speed, refrigerated centrifuge as detailed in section A4 below. |  |  |
| 1.3. | Serviceable Life |  |  |
| 1.3.1. | The System shall have a serviceable life of not less than five years from its Final Acceptance Date (“Serviceable Life”). |  |  |
| 1.4. | The Supplier shall be responsible for the provision of the implementation services for the System as stipulated in section B below. |  |  |
| 1.5. | The Supplier shall be responsible for the provision of the training as stipulated in section C below. |  |  |
| 1.6. | The Supplier shall be responsible for the supply of the documentation for the System as stipulated in section D below. |  |  |
| 1.7. | The Supplier shall be responsible for the performance of acceptance tests as stipulated in section E below. |  |  |
|  | Item no. 1: Benchtop, High-speed Centrifuge |  |  |  |
|  | Intended Use of the System |  |  |
| 2.1.1. | The System is designed for the rapid separation of microsamples, such as blood, urine and other fluidic specimens into their component parts for further diagnostic testing and analysis. |  |  |
|  | Construction and Dimensions |  |  |
| 2.2.1. | The Centrifuge shall have choice of not less than 4 optional fix-angle rotors. |  |  |
| 2.2.2. | The Centrifuge shall be constructed with brushless maintenance free inductive drive. |  |  |
| 2.2.3. | The Centrifuge shall have a dimension (Width x Depth x Height) not exceeding 250 x 360 x 250 mm. |  |  |
| 2.2.4. | The Centrifuge shall have a speed calibration window in the door to record actual rotor speed for calibration purposes. |  |  |
| 2.2.5 | The operator panel, centrifuge lid, chamber and work surface shall be made of materials that are easily cleaned and disinfected. |  |  |
|  | Features, Performance and Functions |  |  |
| 2.3.1. | Rotor and Chamber |  |  |
| 2.3.1.1. | The first Centrifuge shall include a 24 x 1.5/2.0 mL micro-tube rotor. |  |  |
| 2.3.1.2. | The second Centrifuge shall include a 10 x 5 mL tube rotor. |  |  |
| 2.3.1.3. | Each rotor shall come with an aerosol tight with ClickSeal biocontainment lid to ensure the lid is easily and tightly locked to the rotor. |  |  |
| 2.3.1.4. | The maximum RCF (g-force) of the rotors shall be at least 21,000 x g. |  |  |
| 2.3.1.5. | The Centrifuge host shall be able to range maximum speed not less than 14,500 rpm. |  |  |
| 2.3.1.6. | The rotor shall be made of high-performance engineered polymer, which has outstanding chemical resistance to minimise corrosion from aggressive reagents. |  |  |
| 2.3.1.7. | The rotor lid shall be made of transparent material (non-metallic and PC material), allowing the user to inspect the sample tube's condition without opening the lid so as to protect the containment leak out from the rotor chamber. |  |  |
| 2.3.1.8. | A chemical-resistant plastic chamber with up to 48 mL of spilled liquid is retained in the chamber and cannot enter the instrument. |  |  |
| 2.3.1.9. | The centrifuge shall have choice of not less than 4 optional fix-angle rotors. |  |  |
| 2.3.1.10. | A rotor for 24 pieces Hematocrit (capillaries) rotor shall be available for achieving faster capillary separations if requested in the future. |  |  |
| 2.3.2. | Refrigerating Performance and Temperature Control |  |  |
| 2.3.2.1. | Built-in ventilator shall be available. |  |  |
| 2.3.2.2. | Refrigerator shall not be required. |  |  |
| 2.3.3. | Control Panel |  |  |
| 2.3.3.1. | The keypad shall be soft-touch. |  |  |
| 2.3.3.2. | Both keypad and display surface shall be easy to clean. |  |  |
| 2.3.3.3. | The control panel shall be controlled by microprocessor. |  |  |
| 2.3.3.4. | An internal memory to save the most recent data shall be available. |  |  |
| 2.3.3.5. | The control panel shall process the functions for time selection, centrifugation speed, open lid, pulse-spin, start and stop functions, and alarm functions |  |  |
| 2.3.4. | Time Selection |  |  |
| 2.3.4.1. | Run time shall be selected as a run time between 1 and 99 minutes, with a 1-minute increment or continuous operation. |  |  |
| 2.3.4.2. | When the remaining run time is less than 1 minute, the remaining time shall be given in seconds. |  |  |
| 2.3.4.3. | In continuous operation, the time display counts up. The displayed run time is first in seconds. After one minute, the displays change every minute. |  |  |
| 2.3.4.4. | Pulse-spinning shall be available for short-term spinning. The Centrifuge accelerates and brakes at maximum power. |  |  |
| 2.3.4.5. | The displayed run time of pulse-spinning is first in seconds. After one minute the displays changes every minute. |  |  |
| 2.3.5. | Speed Performance |  |  |
| 2.3.5.1. | The minimum speed shall be at least 300 rpm. |  |  |
| 2.3.5.2. | The maximum speed shall not be less than RCF 21,000 x g. |  |  |
| 2.3.5.3. | RPM and RCF reading shall be switchable. |  |  |
| 2.3.5.4. | The acceleration rate shall be 10 seconds or faster, and the deceleration rate shall be 12 seconds or faster. |  |  |
|  | Operation, Power, Electricity and Safety Requirements |  |  |
| 2.4.1. | Operating Condition Requirement |  |  |
| 2.4.1.1. | Environmental conditions: maximum relative humidity 80% up to 31°C; decreasing linearly up to 50% relative humidity at 40°C. |  |  |
| 2.4.1.2. | Permissible ambient temperature during operation: + 5°C to + 40°C. |  |  |
| 2.4.2. | Safety, Certification and Technical Standards |  |  |
| 2.4.2.1. | The Centrifuge lid shall be locked during the centrifugation process. |  |  |
| 2.4.2.2. | The Centrifuge shall equip with automatic imbalance shutdown system. |  |  |
| 2.4.2.3. | The emergency lid release shall be available for mechanical override the door lock during a power failure. |  |  |
| 2.4.2.4. | All rotors shall be autoclavable (120°C, 15 minutes). |  |  |
| 2.4.2.5. | The Centrifuge shall comply safety standards set by IEC 61010-1, IEC 61010-2, IEC 61010-020 or equivalent. |  |  |
| 2.4.2.6. | Aerosol tightness shall be certified by external party such as CAMR, HPA, Porton Down, United Kingdom or equivalent |  |  |
| 2.4.2.7. | The Centrifuge shall be CSA certified, CE marked, IVD compliant, Certified biocontainment. |  |  |
| 2.4.3. | Electricity Supply and Environmental Conditions |  |  |
| 2.4.3.1. | The Centrifuge shall be suitable for operation on local electricity supply of 220 V +/- 6%, 50 Hz +/- 2%, single phase, A.C. |  |  |
|  | Item no. 2: Benchtop, Low-speed, Non-refrigerated Centrifuge |  |  |
| 3.1. | Intended Use of the System |  |  |
| 3.1.1. | The System is designed for the rapid separation of collection tube for blood, urine and other fluidic specimens into their component parts for further diagnostic testing and analysis. |  |  |
| 3.2. | Dimension and Features |  |  |
| 3.2.1. | The Centrifuge shall be bench-top with: |  |  |
| 3.2.1.1. | Width: 560 mm ± 5%. |  |  |
| 3.2.1.2. | Depth: shall be ≤ 700 mm. |  |  |
| 3.2.1.3. | Height: 860 mm ± 5% (open cover). |  |  |
| 3.2.2. | The Centrifuge (without rotor) weight shall be ≤ 95 kg. |  |  |
| 3.2.3 | The Centrifuge shall have choice of at least 8 swing-out rotor configurations and various fix-angle rotors to meet current and future sample processing needs of the lab. |  |  |
| 3.2.4. | The Centrifuge shall have the flexibility for spinning tubes and bottles varying in sizes from 1.5 mL up to 1000 mL when the proper rotor and bucket are used. It shall also be able to spin tissue culture flasks and microplates upon using proper rotors and adapters. |  |  |
| 3.2.5. | The Centrifuge shall be able to accommodate rotors with the capacities as follows: |  |  |
| 3.2.5.1. | 4 x 1000 mL (swinging bucket rotor). |  |  |
| 3.2.5.2. | 4 x 750 mL (swinging bucket rotor). |  |  |
| 3.2.5.3. | 4 x 250 mL (windshield swinging bucket rotor). |  |  |
| 3.2.5.4. | 24 microplates (swinging bucket rotor). |  |  |
| 3.2.5.5. | 20 microplates (swinging bucket rotor). |  |  |
| 3.2.5.6. | 16 microplates (swinging bucket rotor). |  |  |
| 3.2.5.7. | 10 microplates (windshield swinging bucket rotor). |  |  |
| 3.2.5.8. | 6 microplates (swinging bucket rotor). |  |  |
| 3.2.5.9. | 6 x 250 mL (fixed angle). |  |  |
| 3.2.5.10. | 6 x 100 mL (fixed angle). |  |  |
| 3.2.5.11. | 14 x 50 mL (fixed angle). |  |  |
| 3.2.5.12. | 8 x 50 mL (fixed angle). |  |  |
| 3.2.5.13. | 48 x 1.5/2.0 mL (fixed angle). |  |  |
| 3.2.5.14. | 30 x 1.5/2.0 mL (fixed angle). |  |  |
| 3.2.6 | The Centrifuge shall be constructed with brushless maintenance free inductive drive. |  |  |
| 3.2.7. | The material of rotors, buckets, external surface and internal surface shall be chemical resistance to chemical substances and easy to clean by common disinfectant. |  |  |
| 3.3. | Construction, Performance and Functions |  |  |
| 3.3.1. | The Centrifuge drive shall be made of maintenance-free brushless direct inductive drive. |  |  |
| 3.3.2. | Both the housing and the lid shall be made of galvanized steel or equivalent, which are heavy gauge, corrosion and rust resistant. |  |  |
| 3.3.3. | The Centrifuge front shall be made of high-impact resistant plastic. |  |  |
| 3.3.4. | The Centrifuge shall have a speed calibration window in the door to record actual rotor speed for calibration purposes. |  |  |
| 3.3.5. | The operator panel, centrifuge lid, chamber and work surface shall be made of materials that are easily cleaned and disinfected. |  |  |
| 3.3.6. | Control Panel |  |  |
| 3.3.6.1. | The Centrifuge shall come with full colour touch screen interface in which the surface shall be easily cleaned. |  |  |
| 3.3.6.2. | The Centrifuge shall be equipped with microprocessor to control the operation with digital display. |  |  |
| 3.3.6.3. | There shall be display of “set” and “actual run” of speed and time. |  |  |
| 3.3.6.4. | The "set" and the "actual" run conditions shall be easily assessable in the display during operation. |  |  |
| 3.3.6.5. | There shall be an internal memory to save the most recent data. |  |  |
| 3.3.6.6. | The control panel shall process the functions for time and temperature selection, centrifugation speed, open lid, pulse-spin, start and stop functions, and alarm functions. |  |  |
| 3.3.6.7. | The Centrifuge shall be able to display messages in at least seven (7) languages. |  |  |
| 3.3.6.8. | The display shall include at least the following: - Programme identification number; - Speed in rpm and RCF; - Time; - Temperature; - Acceleration rate; - Deceleration rate; - Alert signals or error message to users. |  |  |
| 3.3.7. | Time Selection |  |  |
| 3.3.7.1. | Time shall be adjustable up to not less than 9 hours 59 minutes, in 1-minute increments. |  |  |
| 3.3.7.2. | There shall be a “hold” function for continuous run. |  |  |
| 3.3.7.3. | There shall be two options allowing the timer start counting down at the beginning of the run or at speed to ensure reproducible results under the most challenging conditions. |  |  |
| 3.3.7.4. | While the centrifuge is in operation, the total running time and the running time after reaching the maximum speed can be displayed. |  |  |
| 3.3.7.5. | When the remaining run time is less than 1 minute, the remaining time shall be given in seconds. |  |  |
| 3.3.7.6. | Quick Run Mode: Pressing ‘Quick Run” key activates maximum acceleration up to the maximum permissible speed of rotor, upon key release centrifuge stops with maximum deceleration power. |  |  |
| 3.3.8. | Braking System |  |  |
| 3.3.8.1. | There shall be electronic dynamic on/off and soft brake available. |  |  |
| 3.3.8.2. | Acceleration and Deceleration control: not less than 9 acceleration and 10 deceleration profile choice improves the resolution and yield of separations. |  |  |
| 3.3.8.3. | There shall be two options: allowing the timer to start counting down at the beginning of the run or at speed to ensure reproducible results under the most challenging conditions. |  |  |
| 3.3.8.4. | There shall be an ACE (Accumulated Centrifugal Effect) mode, which is an integrator function that calculates the effect of speed in relation to time and adjusts run time to account for differences in acceleration, helps to transfer applications and their parameter settings between centrifuges. ACE shall ensure that the application runs in exactly the same way and yields the same results as on a legacy centrifuge. |  |  |
| 3.3.9. | Rotor System |  |  |
| 3.3.9.1. | The Centrifuge shall come with a 4 x 1000 mL swing-out rotor with maximum speed at least 3,800 rpm or RCF 3,374 x g. |  |  |
| 3.3.9.2. | Four (4) pieces of 1000 mL round buckets shall be provided. |  |  |
| 3.3.9.3. | Four (4) pieces of aerosol-tight caps shall be provided. |  |  |
| 3.3.9.4. | The aerosol-tight caps shall be certified by Biocontainment certification by Public Health England, Porton Down, United Kingdom or equivalent. |  |  |
| 3.3.9.5. | The aerosol-tight caps shall be easily sealed by flapping lock lever, in which operate in a safe manner without spring clips or metal components, to avoid damage of finger joint when frequently use. |  |  |
| 3.3.9.6. | The aerosol-tight caps shall be easily sealed by flapping lock lever, in which operate in a safe manner without spring clips or metal components, to avoid damage of finger joint when frequently use. |  |  |
| 3.3.9.7. | Transparent for efficient sealing and visual inspection of the content within the bucket. |  |  |
| 3.3.9.8. | Both buckets and caps shall be able to stand with general laboratory disinfectant for decontamination process. |  |  |
| 3.3.9.9. | Adaptors and tube capacity shall meet the following: |  |  |
|  | - 4 pieces of 4.5/6mL blood collection tube adapters shall be provided. The adapters shall be able to hold at least 1664 pieces of 5/7 mL blood tubes per spin; - 4 pieces of 15 mL conical tube adapters shall be provided. The adapters shall be able to hold at least 96 pieces of 15 ml conical tubes per spin; and - 4 pieces of 50 mL conical tube adapters shall be provided. The adapters shall be able to hold at least 40 pieces of 50 mL conical tubes per spin. |  |  |
| 3.3.10. | Speed Performance |  |  |
| 3.3.10.1. | Speed range: from 300 rpm to 15,000 rpm or wider, adjustable in 10 rpm increments or finer. |  |  |
| 3.3.10.2. | Maximum speed: not less than 15,000 rpm and RCF not less than 25,800 x g. |  |  |
| 3.3.10.3. | Maximum Speed (with swing bucket rotor with 750 mL buckets): at speed at least 4,700 rpm or RCF4,800 x g. |  |  |
| 3.3.10.4. | Maximum Speed (with swing bucket rotor with 1,000 mL buckets): at speed at least 3,800 rpm or RCF3,374 x g. |  |  |
| 3.3.10.5. | Maximum speed (with windshield swing out rotor): not less than 6000 rpm and RCF not less than 7100 x g. |  |  |
| 3.3.10.6. | Maximum Speed (with 14 pieces of 50 mL conical tube fix-angle rotor): 10,000 rpm; RCF17,100 x g. |  |  |
| 3.3.10.7. | Maximum Speed (with 8 pieces of 50 mL conical tube fix-angle rotor): 14,500 rpm / RCF 24,000 x g. |  |  |
| 3.3.10.8. | Maximum Speed (with microtube rotor): 15,200 rpm / RCF 25,800 x g. |  |  |
| 3.3.10.9. | For multi-user needs, the operation controls shall be easy to use with both “set” and “actual” run conditions and be easily accessible during operation. The Centrifuge shall also be capable of being set up using rpm or RCF values. |  |  |
| 3.3.10.10. | RCF Determination: Standard for any rotor speed selected. |  |  |
| 3.3.11. | Programming and Security |  |  |
| 3.3.11.1. | The Centrifuge shall have the capability of password protection for: - Prevention of accidental adjustments or removal of programmes; - Ensures samples shall only be removed from the temperature-controlled chamber by the operator who has set the password. |  |  |
| 3.3.11.2. | The Centrifuge shall have capability of password protection for lid opening. |  |  |
| 3.3.11.3. | Time, speed, temperature, acceleration and deceleration profiles have to be saved in the different programmes. |  |  |
| 3.3.11.4. | The centrifuge shall have a minimum of 5 direct recall programme keys for programme 1 to 5, and 1 directory key to load any 1 to 99 programmes by means of the touch assistive control panel. |  |  |
| 3.3.11.5. | The Centrifuge shall have at least 99 user-defined programmes for operation |  |  |
| 3.3.11.6. | The Centrifuge shall be access to the Software ID and Cycle Counts for service maintenance purposes. |  |  |
| 3.4 | Operation, Power, Electricial and Safety Requirements |  |  |
| 3.4.1. | Operation and Safety |  |  |
| 3.4.1.1. | The Centrifuge shall be equipped with automatic rotor recognition. |  |  |
| 3.4.1.2. | The Centrifuge shall have a chamber door controlled by an electromechanical interlock system. |  |  |
| 3.4.1.3. | Soft touch lid lock: There shall be motor assisted lid locking system which improves closure time and reduces wear. |  |  |
| 3.4.1.4. | The Centrifuge shall have an advanced imbalance technology that adapts to specific rotors. |  |  |
| 3.4.1.5. | Dynamic electronic imbalance recognition shall be available in the Centrifuge to monitor rotor load during the total run period. The automatic imbalance recognition shall be optimised for different rotors at different speeds. |  |  |
| 3.4.1.6. | Imbalance tolerance: Imbalance tolerance of the Centrifuge shall not be less than 30 grammes when suited with a 4 x 1000 mL swinging bucket rotor and not less than 20 grammes when suited with a 4 x 750 mL swinging bucket rotor. |  |  |
| 3.4.1.7. | Imbalance Protection: Automatic imbalance shutdown shall be provided in a situation of serious imbalance incidence. |  |  |
| 3.4.1.8. | There shall have a diagnostic display for different system malfunction conditions. |  |  |
| 3.4.1.9. | The Centrifuge shall be equipped with a rotor automatic rotor lock system and automatic rotor recognition for ease of use, time saving and guaranteeing rotor lock security to eliminate the effects of user error and spin error. |  |  |
| 3.4.1.10. | The Centrifuge shall have an option for automatic lid opening at the end of the run. |  |  |
| 3.4.1.11. | The Centrifuge lid shall be supported by gas lid stay to keep the lid stay up when chamber is open. |  |  |
| 3.4.1.12. | An audible alarm at the end of the run shall be available to notify the user. |  |  |
| 3.4.1.13. | The rotor shall be installed and removed with no tool in less than 5 seconds. |  |  |
| 3.4.1.14. | The Centrifuge lid shall be locked during the centrifugation process. |  |  |
| 3.4.1.15. | An emergency lid release shall be available for mechanical override the door lock during a power failure. |  |  |
| 3.4.1.16. | All rotors shall be autoclavable (120°C, 15 minutes) for disinfection. |  |  |
| 3.4.2. | Operation Conditions Requirement |  |  |
| 3.4.2.1. | Environmental conditions: maximum relative humidity 80% up to 31°C; decreasing linearly up to 50% relative humidity at 40°C. |  |  |
| 3.4.2.2. | Permissible ambient temperature during operation+2°C to +35°C. |  |  |
| 3.4.3. | Certification and Technical Standards |  |  |
| 3.4.3.1. | Aerosol tightness shall be certified by external party such as CAMR, HPA, Porton Down, United Kingdom or equivalent. |  |  |
| 3.4.3.2. | The Centrifuge shall be CSA certified, CE marked, IVD compliant, MDD compliant, and Certified biocontainment. |  |  |
| 3.4.3.3. | The Centrifuge shall remain operational and within specification throughout the voltage range of 220V +/- 6%, 50Hz +/- 2% and 1-phase A.C. electrical supply with BS 1363 plug fitted. |  |  |
| 3.4.3.4. | The Centrifuge shall comply with the electromagnetic compatibility (EMC) requirements of IEC 60601-1-2 / IEC 61326-1 or equivalent. |  |  |
| 3.4.3.5. | The manufacturer shall comply with international standard: ISO13485 and ISO9000 or equivalent for quality assurance |  |  |
| 3.4.4. | Electricity Supply and Environmental Conditions |  |  |
| 3.4.4.1. | The Centrifuge shall be suitable for operation on local electricity supply of 220V +/- 6%, 50 Hz +/- 2%, single phase, A.C. |  |  |
| 3.4.4.2. | The noise level at maximum speed should be less than 65 dB(A). |  |  |
|  | Item 3: Centrifuge, Tabletop, Low-Speed, Refrigerated |  |  |  |
|  | Intended Use of the System |  |  |
| 4.1.1. | The System is designed for the rapid separation of collection tube for blood, urine and other fluidic specimens into their component parts for further diagnostic testing and analysis. |  |  |
|  | Dimension and Features |  |  |
| 4.2.1 | The Centrifuge shall be bench-top with: |  |  |
| 4.2.1.1. | Width: 740 mm ± 5% |  |  |
| 4.2.1.2. | Depth: must be less than 700 mm |  |  |
| 4.2.1.3. | Height: 860 mm ± 5% (open cover) |  |  |
| 4.2.2. | The Centrifuge shall have choice of at least 8 swing-out rotor configurations and various fix-angle rotors to meet current and future sample processing needs of the laboratory. |  |  |
| 4.2.3. | The Centrifuge shall have the flexibility for spinning tubes and bottles varying in sizes from 1.5 mL up to 1000 mL when the proper rotor and bucket are used. It shall also be able to spin tissue culture flasks and microplates upon using proper rotors and adapters. |  |  |
| 4.2.4 | The Centrifuge shall be able to accommodate rotors with the capacities as follows: |  |  |
| 4.2.4.1. | 4 x 1000 mL (swinging bucket rotor). |  |  |
| 4.2.4.2. | 4 x 750 mL (swinging bucket rotor). |  |  |
| 4.2.4.3. | 4 x 250 mL (windshield swinging bucket rotor). |  |  |
| 4.2.4.4. | 24 microplates (swinging bucket rotor). |  |  |
| 4.2.4.5. | 20 microplates (swinging bucket rotor). |  |  |
| 4.2.4.6. | 16 microplates (swinging bucket rotor). |  |  |
| 4.2.4.7. | 10 microplates (windshield swinging bucket rotor). |  |  |
| 4.2.4.8. | 6 microplates (swinging bucket rotor). |  |  |
| 4.2.4.9. | 6 x 250 mL (fixed angle). |  |  |
| 4.2.4.10. | 6 x 100 mL (fixed angle). |  |  |
| 4.2.4.11. | 14 x 50 mL (fixed angle). |  |  |
| 4.2.4.12. | 8 x 50 mL (fixed angle). |  |  |
| 4.2.4.13. | 48 x 1.5/2.0 mL (fixed angle). |  |  |
| 4.2.4.14. | 30 x 1.5/2.0 mL (fixed angle). |  |  |
| 4.2.5. | The material of rotors, buckets, external surface and internal surface shall be chemical resistance to chemical substances and easy to clean by common disinfectant. |  |  |
| 4.2.6. | The operator panel, centrifuge lid, chamber and work surface shall be made of materials that are easily cleaned and disinfected. |  |  |
|  | Construction, Performance and Functions |  |  |
| 4.3.1. | The Centrifuge shall be constructed with brushless maintenance free direct inductive drive. |  |  |
| 4.3.2. | Both the housing and the lid shall be made of galvanized steel or equivalent, which are heavy gauge, corrosion and rust resistant. |  |  |
| 4.3.3. | The Centrifuge front is made of high-impact resistant plastic. |  |  |
| 4.3.4. | The Centrifuge shall have a speed calibration window in the door to record actual rotor speed for calibration purposes. |  |  |
| 4.3.5. | Refrigerating performance and Temperature Control |  |  |
| 4.3.5.1. | Temperature control shall be user-programmable from -10°C to +40°C or wider. |  |  |
| 4.3.5.2. | Temperature control is programmable at ≤ 1°C increment. |  |  |
| 4.3.5.3. | The interior of the centrifuge should be maintained at ± 2°C of the programmed temperature during full load at full speed centrifugation. |  |  |
| 4.3.5.4. | The refrigerant shall be CFC-free. |  |  |
| 4.3.5.5. | The refrigerating system shall follow No “Controlled Refrigerant”, as defined under the Ozone Layer Protection (Controlled Refrigerants) Regulation, Cap. 403 of Laws of Hong Kong. |  |  |
| 4.3.5.6. | Pre-temperature function shall be available to allow quick pre-temperature regulation, that is pre-cooling and pre-warming with a range from -10°C to +40°C, of the empty rotor to increase operation efficiency. |  |  |
| 4.3.5.7. | The Centrifuge must be able to display both air/chamber temperature as well as temperature in the sample. |  |  |
| 4.3.6. | Control Panel |  |  |
| 4.3.6.1. | The Centrifuge shall come with full colour touch screen interface in which the surface shall be easy to be cleaned. |  |  |
| 4.3.6.2. | The Centrifuge shall be equipped with microprocessor to control the operation with digital display. Display of set and actual run of speed and time. |  |  |
| 4.3.6.3. | There shall be an internal memory to save the most recent data. |  |  |
| 4.3.6.4. | The control panel shall process the functions for time and temperature selection, centrifugation speed, open lid, pulse-spin, start and stop functions, and alarm functions. |  |  |
| 4.3.6.5. | The Centrifuge shall be able to display messages in at least seven (7) languages. |  |  |
| 4.3.6.6. | Both the "set" and the "actual" run conditions shall be easily assessable in the display during operation. |  |  |
| 4.3.6.7. | The display shall include at least: - Programme identification number; - Speed in rpm and RCF; - Time; - Temperature; - Acceleration rate; - Deceleration rate; - Alert signals or error message to users. |  |  |
| 4.3.7. | Time selection |  |  |
| 4.3.7.1. | Time shall be adjustable up to not less than 9 hours 59 minutes, in 1-minute increments. |  |  |
| 4.3.7.2. | There shall be a “hold” function for continuous run. |  |  |
| 4.3.7.3. | There shall be two options allowing the timer start counting down at the beginning of the run or at speed to ensure reproducible results under the most challenging conditions. |  |  |
| 4.3.7.4. | While the centrifuge is in operation, the total running time and the running time after reaching the maximum speed can be displayed. |  |  |
| 4.3.7.5. | When the remaining run time is less than 1 minute, the remaining time shall be given in seconds. |  |  |
| 4.3.7.6. | Quick Run Mode: Pressing ‘Quick Run” key activates maximum acceleration up to the maximum permissible speed of rotor, upon key release centrifuge stops with maximum deceleration power. |  |  |
| 4.3.8. | Braking System |  |  |
| 4.3.8.1. | There shall be electronic dynamic on/off and soft brake available. |  |  |
| 4.3.8.2. | Acceleration and Deceleration control: not less than 9 acceleration and 10 deceleration profile choice improves the resolution and yield of separations. |  |  |
| 4.3.8.3. | There shall be two options: allowing the timer to start counting down at the beginning of the run or at speed to ensure reproducible results under the most challenging conditions. |  |  |
| 4.3.8.4. | There shall be an ACE (Accumulated Centrifugal Effect) mode, which is an integrator function that calculates the effect of speed in relation to time and adjusts run time to account for differences in acceleration, helps you to transfer applications and their parameter settings between centrifuges. ACE ensures that the application runs in exactly the same way and yields the same results as on a legacy centrifuge. |  |  |
| 4.3.9. | Rotor system |  |  |
| 4.3.9.1. | The Centrifuge shall come with a 4 x 1000 mL swing-out rotor with maximum speed at least 3,800 rpm or RCF 3,374 x g. |  |  |
| 4.3.9.2. | Four (4) pieces of 1000 mL round buckets shall be provided. |  |  |
| 4.3.9.3. | Four (4) pieces of aerosol-tight caps shall be provided. |  |  |
| 4.3.9.4. | The aerosol-tight caps, which certified by Biocontainment certification by Public Health England, Porton Down, United Kingdom or equivalent, shall be included. |  |  |
| 4.3.9.5. | The cap shall be easily sealed by flapping lock lever, in which operate in a safe manner without spring clips or metal components, to avoid damage of finger joint when frequently use. |  |  |
| 4.3.9.6. | Transparent for efficient sealing and visual inspection of the content within the bucket. |  |  |
| 4.3.9.7. | Both buckets and caps shall be able to stand with general laboratory disinfectant for decontamination process. |  |  |
| 4.3.9.8. | Adaptors & tube capacity: 4 pieces of 4.5/6mL blood collection tube adapters shall be provided. The adapters shall be able to hold at least 1664 pieces of 5/7 mL blood tubes per spin. |  |  |
| 4.3.10. | Speed performance |  |  |
| 4.3.10.1 | Speed range: from 300 rpm to 15,000 rpm or wider, adjustable in 10 rpm increments or finer. |  |  |
| 4.3.10.2. | Maximum speed: not less than 15,000 rpm and RCF not less than 25,800 x g. |  |  |
| 4.3.10.3. | Maximum Speed (with swing bucket rotor with 750 mL buckets): at speed at least 4,700 rpm or RCF4,800 x g. |  |  |
| 4.3.10.4. | Maximum Speed (with swing bucket rotor with 1,000 mL buckets): at speed at least 3,800 rpm or RCF3,374 x g. |  |  |
| 4.3.10.5. | Maximum speed (with windshield swing out rotor): not less than 6000 rpm and RCF not less than 7100 x g. |  |  |
| 4.3.10.6. | Maximum Speed (with 14 pieces of 50 mL conical tube fix-angle rotor): 10,000 rpm; RCF17,100 x g. |  |  |
| 4.3.10.7. | Maximum Speed (with 8 pieces of 50 mL conical tube fix-angle rotor): 14,500 rpm / RCF 24,000 x g. |  |  |
| 4.3.10.8. | Maximum Speed (with microtube rotor): 15,200 rpm / RCF 25,800 x g. |  |  |
| 4.3.10.9. | For multi-user needs, the operation controls shall be easy to use with both “set” and “actual” run conditions and be easily accessible during operation. The Centrifuge shall also be capable of being set up using rpm or RCF values. |  |  |
| 4.3.10.10. | RCF Determination: Standard for any rotor speed selected. |  |  |
| 4.3.11. | Programming and Security |  |  |
| 4.3.11.1. | The Centrifuge shall have the capability of password protection for: - Prevention of accidental adjustments or removal of programmes; - Ensures samples are only removed from the temperature-controlled chamber by the operator who has set the password. |  |  |
| 4.3.11.2. | The Centrifuge shall have capability of password protection for lid opening. |  |  |
| 4.3.11.3. | Time, speed, temperature, acceleration and deceleration profiles have to be saved in the different programmes. |  |  |
| 4.3.11.4. | The Centrifuge shall have a minimum of 5 direct recall programme keys for programme 1 to 5, and 1 directory key to load any 1 to 99 programmes by means of the touch assistive control panel. |  |  |
| 4.3.11.5. | The Centrifuge shall have at least 99 user-defined programmes for operation. |  |  |
| 4.3.11.6. | The Centrifuge shall be access to the Software ID and Cycle Counts for service maintenance purposes. |  |  |
|  | Operation, Power, Electricial and Safety Requirements |  |  |
| 4.4.1 | Operation and Safety |  |  |
| 4.4.1.1. | The Centrifuge shall be equipped with automatic rotor recognition. |  |  |
| 4.4.1.2. | The Centrifuge shall have a chamber door controlled by an electromechanical interlock system. |  |  |
| 4.4.1.3. | Soft touch lid lock: shall be motor assisted lid locking system which improves closure time and reduces wear. |  |  |
| 4.4.1.4. | The Centrifuge should have advanced imbalance technology that adapts to specific rotors. |  |  |
| 4.4.1.5. | Dynamic electronic imbalance recognition shall be available in the centrifuge to monitor rotor load during the total run period. The automatic imbalance recognition shall be optimized for different rotors at different speeds. |  |  |
| 4.4.1.6. | Imbalance tolerance: Imbalance tolerance of the centrifuge shall not be less than 30 grammes when suited with a 4 x 1000 mL swinging bucket rotor and not less than 20 grammes when suited with a 4 x 750 mL swinging bucket rotor. |  |  |
| 4.4.1.7. | Imbalance Protection: Automatic imbalance shutdown shall be provided in a situation of serious imbalance incidence. |  |  |
| 4.4.1.8. | The Centrifuge shall have a diagnostic display for different system malfunction conditions. |  |  |
| 4.4.1.9. | The Centrifuge shall be equipped with a rotor automatic rotor lock system and automatic rotor recognition for ease of use, time saving and guaranteeing rotor lock security to eliminate the effects of user error and spin error. |  |  |
| 4.4.1.10. | The Centrifuge shall have an option for automatic lid opening at the end of the run. |  |  |
| 4.4.1.11. | The Centrifuge lid shall be supported by gas lid stay to keep the lid stay up when chamber is open. |  |  |
| 4.4.1.12. | An audible alarm at the end of the run shall be available to notify the user. |  |  |
| 4.4.1.13. | The rotor shall be installed and removed with no tool in less than 5 seconds. |  |  |
| 4.4.1.14. | The Centrifuge lid shall be locked during the centrifugation process. |  |  |
| 4.4.1.15. | An emergency lid release shall be available for mechanical override the door lock during a power failure. |  |  |
| 4.4.1.16. | All rotors shall be autoclavable (120°C, 15 minutes) for disinfection. |  |  |
| 4.4.2. | Operation conditions requirement |  |  |
| 4.4.2.1. | Environmental conditions: maximum relative humidity 80% up to 31°C; decreasing linearly up to 50% relative humidity at 40°C. |  |  |
| 4.4.2.2. | Permissible ambient temperature during operation +2°C to +35°C. |  |  |
| 4.4.3. | Certification and Technical Standards |  |  |
| 4.4.3.1. | Aerosol tightness shall be certified by external party such as CAMR, HPA, Porton Down, U.K or equivalent. |  |  |
| 4.4.3.2. | The instrument shall be CSA certified, CE marked, IVD compliant, MDD compliant, and Certified biocontainment. |  |  |
| 4.4.3.3. | The equipment shall remain operational and within specification throughout the voltage range of 220 V +/- 6%, 50 Hz +/- 2% and 1-phase A.C. electrical supply with BS 1363 plug fitted. |  |  |
| 4.4.3.4. | Equipment offered shall comply with the electromagnetic compatibility (EMC) requirements of IEC 60601-1-2 / IEC 61326-1 or equivalent. |  |  |
| 4.4.3.5. | The manufacturer shall comply with international standard: ISO13485 and ISO9000 or equivalent for quality assurance. |  |  |
| 4.4.4. | Electricity Supply and Environmental Conditions |  |  |
| 4.4.4.1. | The instrument shall be suitable for operation on local electricity supply of 220V +/- 6%, 50 Hz +/- 2%, single phase, A.C. |  |  |
| 4.4.4.2. | The noise level at maximum Speed should be less than 65 dB(A). |  |  |
|  | Implementation Services |  |  |
|  | The System shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included within 10 weeks from the date informed by CMHHK. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
|  | Training |  |  |
|  | The Supplier shall be responsible to provide session of on-site maintenance training to representatives of CMHHK upon request at no additional charges. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures, calibration and alignment, adjustment. |  |  |
|  | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. |  |  |
|  | Documentation |  |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |
|  | At the time of delivery of the equipment, appropriate Set(s) of the manufacturer’s original operation and maintenance manuals in English or in Chinese complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to CMHHK after order placement but before the equipment delivery for documentation purpose. |  |  |
|  | The Supplier shall provide the authorization letter from manufacturer for providing products, parts and service to end user. The written undertaking shall be signed by a duly authorized representative of the manufacturer and dated no later than the quotation closing date. |  |  |
|  | Acceptance Tests |  |  |  |
|  | Safety Test For the purpose of this contract the System shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Contractor with the witness of representative of CMHHK. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the System. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
|  | Functional Test For the purpose of this Contract the System shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. In the event that the equipment fails to conform to the above stated requirements, the successful Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the Systems shall be determined by CMHHK based upon the satisfactory completion of such functional test. |  |  |
|  | Indicative Warranty Service |  |  |
|  | The Supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The potential Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
|  | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least one times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |
|  | The Supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure which may arise within the warranty period of the equipment. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the potential Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
|  | Indicative Maintenance Service |  |  |  |
|  | The Supplier shall quote the charge for annual maintenance services after the warranty period within the serviceable life of the proposed System. |  |  |
|  | The Supplier shall submit a price list of all spare parts of the System chargeable to the CMHHK Operator. For spare parts not covered by the submitted prices, the potential Supplier must submit a quotation to the CMHHK Operator for consideration every time when spares are required. |  |  |
|  | The Supplier shall deploy properly trained service personnel to carry out the maintenance services and shall ensure that all necessary precautions for their safety are taken. |  |  |
|  | The Supplier shall provide free of additional charge corrective maintenance service for providing immediate repair service for the System and related equipment in normal working hours. |  |  |
|  | The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the potential Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work is not completed at the end of particular normal working period, subject to the user’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
|  | Spare Parts |  |  |
|  | The Supplier shall guarantee the availability of maintenance spare parts for the anticipated life of the System. Sufficient spare parts shall be held by the successful Supplier to cater for the maintenance during the warranty period. |  |  |
|  | The Suppliers, in their tender submission, shall provide a comprehensive list of recommended spare parts with unit prices valid for at least one (1) year after expiry of warranty. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Centrifuge can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Centrifuge does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Centrifuge in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| IEC61010-1 or equivalent, IEC61010-2 or equivalent, and IEC61010-20 or equivalent | Details please refer to the Section 2.4.2.5. in Part 3. |  |  |  |
| Certified by CAMR, HPA, Porton Down, Unitied Kingdom or equivalent. | Details please refer to the Section 2.4.2.6., 3.4.3.1. & 4.4.3.1 in Part 3. |  |  |  |
| CSA certified, CE marked, IVD compliant, and Certified biocontainment. | Details please refer to the Section 2.4.2.7. in Part 3. |  |  |  |
| CSA certified, CE marked, IVD compliant, MDD compliant, and Certified biocontainment. | Details please refer to the Section 3.4.3.2. & 4.4.3.2. in Part 3. |  |  |  |
| Electromagnetic compatibility (EMC) requirements of IEC 60601-1-2 or equivalent / IEC 61326-1 or equivalent. | Details please refer to the Section 3.4.3.4.& 4.4.3.4. in Part 3. |  |  |  |
| ISO13485 or equivalent, and ISO9000 or equivalent | Details please refer to the Section 3.4.3.5. & 4.4.3.5. in Part 3. |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1-1.7 in Part 3, including the provision of a minimum 12-month warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge* (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

Note: * The Total One-time Charge shall include one-year of warranty period.

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-month period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-month period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END