P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Endoscopic Viewing System

for The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

(CMHPO Ref.           )

To : 	Project Director (CMHPO)

(Attn. Christine YEUNG)

### [by fax: 2127 4795 or email: cklyeung@healthbureau.gov.hk]

Your ref:

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

RFI (CE)

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply of Endoscopic Viewing System (hereinafter refers as the “System”) for The Chinese Medicine Hospital of Hong Kong (“CMHHK”) located at Pak Shing Kok in Tseung Kwan O. The CMHPO therefore wishes to collect market information on this System.

- Background of the CMHHK Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK. CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK. The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018. Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK. The CMHHK project has proceeded to the commissioning stage in 2021. The hospital services have been commenced by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company has more than one Endoscopic Viewing System that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Endoscopic Viewing System.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section G in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ________________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “Goods” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only. They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |
| 1 | Overall Requirements |  |  |
| 1.1 | The Goods shall be supplied for two (2) sets of Endoscopic Video Systems (the “Goods”) comprising the following: |  |  |
|  | - two (2) sets of video processor as detailed in section A2 below; |  |  |
|  | - two (2) sets of LED light source unit (or an integral unit with video processor) as detailed in section A3 below; |  |  |
|  | - two (2) sets of Medical monitor as detailed in section A4 below; |  |  |
|  | - three (3) sets of Video Gastroscope as detailed in section A5 below; |  |  |
|  | - three (3) sets of Video Colonoscope as detailed in section A6 below; |  |  |
|  | - two (2) sets of Video Bronchoscope as detailed in section A7 below; |  |  |
|  | - two (2) sets of Endoscopic CO2 insufflator as detailed in section A8 below; |  |  |
|  | - two (2) sets of irrigating pump as detailed in section A9 below; |  |  |
|  | - two (2) sets of Medical recorder unit as detailed in section A10 below; |  |  |
|  | - two (2) sets of real-time lesion automated detection system (or an integral feature in the video processor) as detailed in section A11 below; |  |  |
|  | - two (2) sets of video printer as detailed in section A12 below; and |  |  |
|  | - two (2) sets of equipment cart with monitor holding arm as detailed in section A13 below. |  |  |
| 1.2 | The offer shall be completed with all necessary accessories and interconnection that are essential for the normal operation of the equipment. Subject to request by hospital, the Supplier might be required to install the offered Goods on surgical pendant shelves. Video signal connections between video monitor on existing monitor arm and pendant equipment; therefore, shall be included and supported by competent party. |  |  |
| 2 | Video Processor |  |  |
| 2.1 | It shall be used for full HD (or better) image reproduction of flexible video endoscopes. |  |  |
| 2.2 | The processor shall be able to capture endoscopic image with lifelike clarity on even minute capillaries. |  |  |
| 2.3 | It shall offer multiple image enhanced endoscopy features for improved visibility. Supporting documentation shall be provided for evaluation. |  |  |
| 2.4 | Scanning Method: Simultaneous |  |  |
| 2.5 | Scanning resolution: 1080p or better |  |  |
| 2.6 | Image Enhanced Endoscopy features |  |  |
|  | (a) Light spectrum limited imaging technique (1): the mode shall enhance the visibility of capillaries and other structures on the mucosal surface. |  |  |
|  | (b) Light spectrum limited imaging technique (II): the second mode shall increase the mucosal contrast with the use of light source spectrum different than (a) technology employing additional wavelength (e.g. amber light) to distinguish different blood vessels is also acceptable. |  |  |
|  | (c) Digital processing image enhancement function: this mode shall enhance observation of tissue texture and color as well as blood vessels by using digital signal processing to extract the spectral imaged of specific wavelengths from white light. |  |  |
| 2.7 | Captured images shall be displayed at least four (4) images during procedure for better observation. |  |  |
| 2.8 | Contrast: shall be adjustable at least three levels |  |  |
| 2.9 | Color Adjustment: shall be brightness, Red, Blue, Chroma at least nine levels |  |  |
| 2.10 | White balance free: shall be equipped to minimize the set-up effort when insertion of the endoscope or exchange the endoscope during the diagnosis. |  |  |
| 2.11 | Automatic gain control (AGC) or equivalent function: the image shall be electrically amplified when the light is inadequate due to the distal end of the endoscope being too far from the object. |  |  |
| 2.12 | Iris mode: selectable |  |  |
|  | (a) Average: control brightness in general |  |  |
|  | (b) Peak: control brightness in highlight areas |  |  |
|  | (c) Auto |  |  |
| 2.13 | Image enhancement setting: fine patterns or edges in the endoscopic image shall be enhanced electronically to increase the image sharpness and structure. |  |  |
| 2.14 | Tone mode function or equivalent: shall be emphasized slight differences between colors by emphasizing the degree of vividness of color. |  |  |
| 2.15 | Peak detection function for obtaining the highest contrast image shall be available. |  |  |
| 2.16 | The size of the endoscopic image shall be changeable. |  |  |
| 2.17 | Picture in Picture: It shall be available for any combination of endoscopic images, ultrasound images and others. |  |  |
|  | (a) At least two (2) digital videos input terminals, allowing for two Picture out Picture video images to be displayed simultaneously on the endoscopy monitor. |  |  |
| 2.18 | Electronic zoom: x1.0 to x1.5 in multiple steps with full focus image. |  |  |
| 2.19 | The following settings shall be reset to their defaults: user pre-set, image source color tone, freeze, release index, zoom, special light observation, stopwatch, character on screen, exposure, Picture in Picture, etc. |  |  |
| 2.20 | Freeze mode: an endoscopic image shall be frozen by using a scope switch. |  |  |
|  | (a) Real time image shall be displayed at the same time during the image is frozen to prevent missing of detection. |  |  |
| 2.21 | Video output |  |  |
|  | (a) Analogue Image Output: DTV OUT (composite or more) |  |  |
|  | (b) Digital Image Output: HDTV OUT (DVI or HD-SDI or more). HDTV output shall be sufficient to medical monitor and peripheral device (e.g. medical recorder) without adding external video splitter. |  |  |
| 2.22 | Documentation |  |  |
|  | (a) The following ancillary equipment shall be controlled (compatible models): video printer, video recorder, etc.. |  |  |
|  | (b) The following data shall be displayed on the monitor by keyboard: patient ID, patient demographics, date of recording, image frame number, display image setting, structure emphasis, tone, electronic zoom ratio, observation modes, etc.. |  |  |
|  | (c) The following scope-related data stored in the memory chip of the scope shall be recalled: scope model, serial number, etc.. |  |  |
| 2.23 | Image capture and playback shall be available. |  |  |
| 2.24 | The offer shall be DICOM compatible |  |  |
| 2.25 | Type of IEC protection classification against electric shock: Class I, Type BF applied part or better. |  |  |
| 3 | LED Light Source Unit |  |  |
| 3.1 | It shall be intended to be used in combination with flexible video endoscope, medical monitor, recorder and related peripheral equipment for observation, diagnosis and endoscopic treatment. |  |  |
| 3.2 | The offer can be an integral part of the video processor or a dedicated light source unit used in conjunction with video processor offered in Clause 2. |  |  |
| 3.3 | It shall automatically control brightness to achieve ideal illumination for observation. |  |  |
| 3.4 | It shall equip with at least four (4) numbers of different LED sources. |  |  |
| 3.5 | White balance free: shall be equipped to minimize the set-up effort when insertion of the endoscope or exchange the endoscope during the diagnosis. |  |  |
| 3.6 | It shall have back light front panel indicator and control improve operability. |  |  |
| 3.7 | It shall have waterproof one-touch connector. |  |  |
| 3.8 | Illumination |  |  |
|  | (a) Illumination source: four (4) numbers of LED or more. |  |  |
|  | (b) Brightness adjustment: automatically and manually. |  |  |
| 3.9 | Cooling: Forced-Air cooling or better |  |  |
| 3.10 | Multiple special light spectrum limited observation mode shall be available to enhance visualization. |  |  |
| 3.11 | Indicator on the front panel. |  |  |
|  | (a) Transmitted illumination for checking the position of the distal end from outside the body |  |  |
|  | (b) Illumination mode: It lights up if special light mode switch on. |  |  |
| 3.12 | Setting memory |  |  |
|  | (a) Set value shall be maintained after the system is off. |  |  |
| 3.13 | Classification as Medical Electrical Equipment |  |  |
|  | (a) Type of Protection against Electric Shock Class I. |  |  |
|  | (b) Degree of protection against electric shock Type BF applied part or better. |  |  |
| 4 | Medical Monitor |  |  |
| 4.1 | It shall be designed to accurately reproduce the enormous amount of image data contained in the HDTV signal equipped with IPS Technology or equivalent. definition imaging capability shall ensure smooth, flicker-free images, while the screen shall provide a large easy-to-see picture. |  |  |
| 4.2 | Screen size: 32 inch or larger with at least 178o viewing angle. The monitor frame shall be slim in design. |  |  |
| 4.3 | Resolution: full HD resolution or better |  |  |
| 4.4 | Contrast ratio: 1000:1 or better |  |  |
| 4.5 | Luminance (Panel Spec.): 450cd/m2(typical) or better |  |  |
| 4.6 | Aspect ratio: 16:9 or better |  |  |
| 4.7 | Number of colors: 1.07 billion or more |  |  |
| 4.8 | Backlight: LED |  |  |
| 4.9 | Input Terminal: |  |  |
|  | (a) DVI/ HDMI/ DP |  |  |
|  | (b) 3G -SDI/12G-SDI |  |  |
| 4.10 | Output Terminal: |  |  |
|  | (a) 3G-SDI/12G-SDI output |  |  |
|  | (b) CLONE Output |  |  |
| 4.11 | The flat surface shall allow user to easily wipe liquids and gels off from the LCD panel and control buttons to facilitate cleanliness and disinfection. |  |  |
| 5 | Video Gastroscope |  |  |
| 5.1 | The image sensor shall be a CMOS/CCD built into the scope tip for Full HD-resolution image quality (or better) and support the advanced imaging features of video processor for improved tissue visualization. |  |  |
| 5.2 | Field of view: 140 or wider |  |  |
| 5.3 | Direction of view: 0 forward viewing |  |  |
| 5.4 | Depth of field: 3 – 100 mm or wider range |  |  |
| 5.5 | Auxiliary water jet function shall keep the view clear by washing off mucous and enabling past pinpointing of bleeding sites. |  |  |
| 5.6 | Distal End Outer diameter: ≤9.9 mm. |  |  |
| 5.7 | Insertion Tube Outer Diameter: <9.8 mm. |  |  |
| 5.8 | Bending Section: |  |  |
|  | (a) Up: 210 or more |  |  |
|  | (b) Down: 90 or more |  |  |
|  | (c) Right: 100 or more |  |  |
|  | (d) Left: 100 or more |  |  |
| 5.9 | Working Length: 1100mm +/- 10% |  |  |
| 5.10 | Instrument channel inner diameter: ≥ 2.8mm |  |  |
| 5.11 | One-step connector shall be equipped to improve durability and hygienicity. |  |  |
| 5.12 | Scope ID function It shall incorporate a memory chip that stores information about that scope such as its model, serial number, and basic specification. When the scope is connected to the video processor, the scope’s information is automatically transmitted to the system. |  |  |
| 5.13 | The air/water valve, suction valve, and cleaning accessories shall be autoclavable that needs no lubrication for thorough cleaning and disinfecting. Single-use item will also be considered. |  |  |
| 5.14 | It shall be fully immersed in disinfecting solution for thorough automated reprocessing to reduce the risk of cross-contamination. |  |  |
| 5.15 | The video gastroscope shall be compatible with the video processor offered with full support of the Image Enhanced Endoscopy features mentioned in Clause 2.6. |  |  |
| 6 | Video Colonoscope |  |  |
| 6.1 | The image sensor shall be a CMOS/CCD built into the scope tip for Full HD-resolution image quality (or better) and support the advanced imaging features of video processor for improved tissue visualization. |  |  |
| 6.2 | Field of view: 170 or wider |  |  |
| 6.3 | Direction of view: 0 forward viewing |  |  |
| 6.4 | Depth of field: 3 – 100 mm or wider range |  |  |
| 6.5 | Auxiliary water jet function shall be equipped to keep the view clear by washing off mucous and enabling past pinpointing of bleeding sites. |  |  |
| 6.6 | Distal End Outer diameter: ≤13.2 mm. |  |  |
| 6.7 | Insertion Tube Outer Diameter: ≤12.8 mm. |  |  |
| 6.8 | Bending Section: |  |  |
|  | (a) Up: 180 or more |  |  |
|  | (b) Down: 180 or more |  |  |
|  | (c) Right: 160 or more |  |  |
|  | (d) Left: 160 or more |  |  |
| 6.9 | Working Length: 1690mm. +/- 10% |  |  |
| 6.10 | Instrument channel inner diameter: ≥ 3.7mm. |  |  |
| 6.11 | One-step connector shall be equipped to improve durability and hygienicity. |  |  |
| 6.12 | Scope ID function It shall incorporate a memory chip that stores information about that scope such as its model, serial number, and basic specification. When the scope is connected to the video processor, the scope’s information is automatically transmitted to the system. |  |  |
| 6.13 | The air/water valve, suction valve, and cleaning accessories shall be autoclavable that needs no lubrication for thorough cleaning and disinfecting. Single-use item will also be considered. |  |  |
| 6.14 | It shall be fully immersed in disinfecting solution for thorough automated reprocessing to reduce the risk of cross-contamination. |  |  |
| 6.15 | The video colonoscope shall be compatible with the video processor offered with full support of the Image Enhanced Endoscopy features mentioned in Clause 2.6 |  |  |
| 7 | Video Bronchoscope |  |  |
| 7.1 | The image sensor shall be a CMOS/CCD built into the scope tip for Full HD-resolution image quality (or better) and support the advanced imaging features of video processor for improved tissue visualization. |  |  |
| 7.2 | Field of view: 120 or wider |  |  |
| 7.3 | Direction of view: 0 forward viewing |  |  |
| 7.4 | Depth of field: 2-100 mm or wider range |  |  |
| 7.5 | Distal End Outer diameter: ≤ 5.5 mm. |  |  |
| 7.6 | Insertion Tube Outer Diameter: ≤ 5.6 mm. |  |  |
| 7.7 | Bending Section: |  |  |
|  | (a) Up: 210 or more |  |  |
|  | (b) Down: 130 or more |  |  |
| 7.8 | Working Length: 600mm +/- 10% |  |  |
| 7.9 | Instrument channel inner diameter or size: ≥ 2.2mm |  |  |
| 7.10 | One-step connector shall be equipped to improve durability and hygienicity. |  |  |
| 7.11 | Scope ID function It shall incorporate a memory chip that stores information about that scope such as its model, serial number, and basic specification. When the scope is connected to the video processor, the scope’s information is automatically transmitted to the system. |  |  |
| 7.12 | The air/water valve, suction valve, and cleaning accessories shall be autoclavable that needs no lubrication for thorough cleaning and disinfecting. Single-use item will also be considered. |  |  |
| 7.13 | It shall be fully immersed in disinfecting solution for thorough automated reprocessing to reduce the risk of cross-contamination. |  |  |
| 7.14 | It shall be able to be sterilized using the plasma sterilization. |  |  |
| 7.15 | The video bronchoscope shall be compatible with the video processor offered with full support of the Image Enhanced Endoscopy features mentioned in Clause 2.6 |  |  |
| 8 | Endoscopic CO2 Insufflator |  |  |
| 8.1 | The endoscopic insufflator offered shall be intended for endoscopic procedure. It shall be fully compatible with the endoscope in offered for CO2 gas. |  |  |
| 8.2 | Reusable Water Container with CO2 gas inlet design and gas tube shall be included. Disposable option shall also be available for consideration. |  |  |
| 8.3 | Insufflation on/off panel switch. |  |  |
| 8.4 | Length of gas tube shall be at least 1000mm.. |  |  |
| 8.5 | The offered insufflator shall be compatible with medical gas pipeline system in the users’ hospital. The medical gas hose shall comply with ISO 5359 standard and the CO2 gas probe shall be complied with BS 5682 style. If the users’ hospital does not consist of central medical CO2 supply, a pin index York connector for CO2 cylinder shall be provided. |  |  |
| 8.6 | The reusable water container shall be fully immersed in disinfecting solution for thorough automated reprocessing to reduce the risk of cross-contamination. |  |  |
| 8.7 | The reusable water container shall able to be sterilized by either steam sterilization or plasma sterilization. |  |  |
| 8.8 | The endoscopic insufflator equipped with supply pressure indicator to alert for CO2 pressure drops. |  |  |
| 9 | Irrigation Pump |  |  |
| 9.1 | The irrigation pump shall be designed to be used with gastrointestinal endoscopes and colonoscopies to facilitate washing of tissues to remove blood, faeces and other organic matter, enabling improved visualization, diagnosis and therapy during endoscopic procedures. The pump shall allow water to be delivered into the GI tract to create a medium in which endoscopic ultrasound examinations can be performed. |  |  |
| 9.2 | It shall be fully compatible with the endoscope. |  |  |
| 9.3 | Flow rate shall be adjustable from front panel. |  |  |
| 9.4 | Water feed controlled by pneumatic footswitch or remotely from endoscope via compatible video processor. |  |  |
| 9.5 | Water container, tubing and filter shall be provided |  |  |
| 10 | Medical Recorder Unit |  |  |
| 10.1 | The offered medical recorder shall be designed for capturing clinical pictures and videos in Full HD standard (or better) throughout the endoscopic procedure without intervening the system control of endoscopic video processor. |  |  |
| 10.2 | The video signal input of recorder shall be directly connected to video endoscopic system without degradation in image quality. |  |  |
| 10.3 | Voice shall be recorded at the same time of video capturing. |  |  |
| 10.4 | Data shall be recorded on both internal hard disk and external media during procedure. A preview screen of recording image/ video shall be available on the console. |  |  |
| 10.5 | Video recording format: - HD:1920x1080p, 1920x1080i, 1280x720p or more |  |  |
| 10.6 | Internal HDD capacity: 1TB or more. Display of remaining recording time shall be available with the selected resolution. |  |  |
| 10.7 | External USB storage media or equivalent shall be supported. |  |  |
| 11 | Real-time lesion automated detection system |  |  |
| 11.1 | The offered computer-aided image detection system shall be designed for automated/semi-automated real-time detection of lesions such as colonic polyps, during colonoscopy examination with compatible endoscopes. |  |  |
| 11.2 | The offer can be an integral feature of the video processor or a dedicated image processing system used in conjunction with video processor offered in clause 2. |  |  |
| 11.3 | The offer shall consist video output port for real-time processed video to external monitor and/or medical recorder (DVI/ HDMI or HD-SDI port). |  |  |
| 11.4 | An intuitive marker or equivalent shall indicate on the real-time processed video to indicate detection of a potential lesion. |  |  |
| 11.5 | The software in the detection system shall be able to update via a USB port, the supplier should maintain the software to the latest version. |  |  |
| 11.6 | Other AI Options If there is any other AI options available beside clause 13, please quote as optional items for hospitals' consideration. |  |  |
| 12 | Video Printer |  |  |
| 12.1 | The Video Color Printer with digitized processing technology to reproduce full color copy images from video signals. |  |  |
| 12.2 | Resolution: 420 dpi or better |  |  |
| 12.3 | Gradation: 256 level (8 bits) gradations for yellow, magenta, cyan or better. |  |  |
| 12.4 | Effective print pixels: L size: 2132 x 1600 dots or better |  |  |
| 12.5 | Print Size: L size: 127.9 x 96mm. |  |  |
| 12.6 | It shall have a front LCD which provides better visibility for intuitive operation. |  |  |
| 13 | Equipment Cart with Monitor Holding Arm |  |  |
| 13.1 | The equipment cart shall be designed for accommodation of the offered equipment including but not limited to video endoscopic system, peripheral equipment of endoscopic system (e.g. insufflator, irrigation pump), medical monitor, keyboard, isolation transformer, etc.. |  |  |
| 13.2 | It shall have at least four (4) castors, with at least two (2) are lockable. The cart shall be able to move and turn by users. |  |  |
| 13.3 | The equipment cart shall contain the electrical wiring and signal cable of endoscopic systems properly. Handles or equivalent ergonomic feature shall be provided on the cart to enable manoeuvrability. |  |  |
| 13.4 | The equipment cart shall consist of at least three (3) shelves and endoscope holder. Isolation transformer shall be included and securely installed on the equipment cart. |  |  |
| 13.5 | The monitor holder shall be equipped with a swivel arm and be able to accommodate a monitor mentioned in Clause 4. |  |  |
| 13.6 | The equipment cart shall have a scope pole with scope hanger shall be able to store to two (2) endoscopes. |  |  |
| 13.7 | The equipment cart shall have a sliding keyboard tray. |  |  |
| 13.8 | A CO2 cylinder holder shall be provided if central CO2 gas supply is not equipped in the endoscopy centre. |  |  |
| 14 | Safety Requirements |  |  |
| 14.1 | Equipment offered and peripheral components shall comply with the safety requirements of IEC60601-1 or equivalent. |  |  |
| 14.2 | Equipment offered and peripheral components shall comply with the electromagnetic compatibility (EMC) requirements of IEC 60601-1-2 or equivalent. |  |  |
| 14.3 | The offered video endoscopic system shall be approved and listed as medical device in national or international authorities (e.g. NMPA/ DH/ FDA/ MDR). |  |  |
| 14.4 | The equipment shall remain operational and within specification throughout the voltage range of 220V ± 6%, 50Hz ± 2%, 1-phase A.C. electrical supply. |  |  |
| 14.5 | Single phase mains operated equipment shall be fitted with a power plug suitable for the site condition. The plug shall comply with relevant standards e.g. BS1363 for 13A plug. |  |  |
| 14.6 | The equipment shall be effectively bonded to earth unless it is double insulated. |  |  |
| 15 | Miscellaneous Requirement |  |  |
| 15.1 | The video endoscopes in the offer shall be compatible with common brand of Automated Endoscope Reprocessor (AER) from third party manufacturer (e.g. Getinge, Medivators, and Wassenburg etc.). Hook-up connector with third party AER shall be separately quoted as optional item for reference. Official certification provided by third party manufacturer shall be submitted with the tender. |  |  |
| 15.2 | The video endoscopes in the offer shall be compatible with common brand of Endoscopic Reporting System. |  |  |
| 15.3 | The video endoscopes in the offer shall be compatible with common brand of Leakage Tester. |  |  |
| 15.4 | The video endoscopes in the offer shall be compatible with common brand Drying and Storage System for endoscope. |  |  |
| 15.5 | The offer shall be completed with all necessary accessories that are essential for the normal operation of the equipment. Details of the offer shall be specified by the Supplier. |  |  |
| 15.6 | The price quoted shall cover the equipment cost, delivery and installation of the whole set of equipment. |  |  |
| B | Implementation Services |  |  |
| 1 | The Supplier shall be responsible for connecting all waste outlets, data cable and electricity supply to the Goods and any modification works if required. |  |  |
| 2 | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
| 3 | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
| 4 | The Supplier shall be responsible for the make good of ceiling, wall and floor in the installation site after completion of equipment installation. |  |  |
| 5 | The Goods delivery, installation and acceptance tests must be completed within 16 weeks from the date informed by the hospital. |  |  |
| C | Training |  |  |
| 1 | On-site maintenance and operational training shall be provided at no additional charges for maintenance and operation staff. |  |  |
| 2 | The Supplier shall be responsible to provide a session of on-site maintenance training to representatives of CMHHK upon request. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures, calibration and alignment, adjustment. |  |  |
| 3 | The time-table and commencement dates for the training shall be advised at least one month prior to the commencement of the course. Detailed syllabuses shall be submitted for approval, upon request. The practical part of the training shall coincide with the installation and commissioning of the Goods. |  |  |
| 4 | Additional on-site application sessions shall also be arranged upon request by users during the warranty period without additional cost. |  |  |
| 5 | All training materials, such as course, video or slide presentation, diagrams and flowcharts, shall be available at the time of training. At least one copy of material, written in English, shall be provided to each trainee. |  |  |
| D | Documentation |  |  |
| 1 | All photocopies of operation and maintenance manuals shall be properly blinded, stamped and certified as true copies of the original by the manufacturer. |  |  |
| 2 | At the time of delivery of the equipment, two set(s) of the manufacturer’s original operation and maintenance manuals in English complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
| 3 | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to the hospital after order placement but before the equipment delivery for documentation purpose. |  |  |
| 4 | The operation manual shall also be available as a standard softcopy PDF format. |  |  |
| E | Acceptance Tests |  |  |
| 1 | Safety Test For the purpose of this contract, the Goods shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Supplier with the witness of representative of CMHHK. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the Goods. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
| 2 | If the Goods fail to comply with the safety standards listed in above clause, the Supplier shall investigate the causes of the failure and take appropriate corrective actions. The hospitals must be immediately informed of the failure. |  |  |
| 3 | In the case of medical equipment is reported with incident or safety problem, the Supplier shall: (a) attend to the call at site as soon as practically reasonable; (b) conduct investigation to identify the cause of the incident causing the safety problem; (c) advise the hospitals on the recommended actions to be taken based on investigation results; (d) submit interim report and full investigation report to the hospitals within a time frame agreed by both parties. |  |  |
| 4 | Functional Test For the purpose of this Contract the Goods shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. If the Goods fail to comply with functional test, Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the System shall be determined by the Hospital based upon the satisfactory completion of such functional test. |  |  |
| 5 | The Supplier shall submit the acceptance test schedule, procedures, forms and testing method to the end-user for prior approval before the tests. |  |  |
| 6 | Acceptance test shall be performed on-site and the test report shall be submitted. All testing equipment shall be provided by Supplier and shall be calibrated. |  |  |
| F | Indicative Warranty Service |  |  |
| 1 | The Supplier shall guarantee the Goods or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
| 2 | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the Goods offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least two times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
| 3 | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday). |  |  |
| 4 | The Supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the Goods due to improper workmanship, faulty design or component failure which may arise within the warranty period of the Goods. |  |  |
| 5 | The Supplier shall guarantee the serviceable life span of the Goods for a period of not less than 8 years, from the date of acceptance of the Goods; otherwise, the Supplier shall guarantee to provide another set of at least similar (but not lower end) grade of Goods to fulfil the 8 years’ service commitment. Supplier shall also guarantee to supply all reagents and spare parts for at least 10 years from that same date. |  |  |
| 6 | Upon notification by the CMHHK Operator of a defect in the operation of the Goods of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the Goods to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
| 7 | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the system breakdown investigation and corrective action taken. |  |  |
| 8 | Manufacturer-trained service engineers stayed locally to provide on-call services shall be available. |  |  |
| G | Indicative Maintenance Service |  |  |
| 1 | The Supplier shall quote the charge for annual maintenance services after the warranty period within the serviceable life of the proposed Goods. |  |  |
| 2 | The Supplier shall submit a price list of all spare parts of the Goods chargeable to the CMHHK Operator. For spare parts not covered by the submitted prices, the Supplier must submit a quotation to the CMHHK Operator for consideration every time when spares are required. |  |  |
| 3 | The Supplier shall deploy properly trained service personnel to carry out the maintenance services and shall ensure that all necessary precautions for their safety are taken. |  |  |
| 4 | The Supplier shall provide free of additional charge corrective maintenance service for providing immediate repair service for the Goods. |  |  |
| 5 | The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
| 6 | Upon notification by the CMHHK Operator of a defect (departure from performance specifications) in the operation of the Goods of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the Goods to its normal operational conditions in a time of no more than 3 working days. If such work is not completed at the end of particular normal working period, subject to the user’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the supplier to carry on working until the particular maintenance task is completed. |  |  |
| 7 | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
| H | Spare Parts |  |  |
| 1 | The supplier shall guarantee the availability of maintenance spare parts for the anticipated life of the Goods. Sufficient spare parts shall be held by the Supplier to cater for the maintenance during the warranty period. |  |  |
| 2 | The Supplier, in their tender submission, shall provide a comprehensive list of recommended spare parts with unit prices valid for at least one (1) year after expiry of warranty. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Endoscopic Viewing System can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Endoscopic Viewing System does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Endoscopic Viewing System in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| ## IEC60601-1 or equivalent | ## As per the requirements stated in clause 14.1 |  |  |  |
| IEC 60601-1-2 or equivalent | As per the requirements stated in clause 14.2 |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section H in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END