P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Automated Endoscopic Reprocessor

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

(CMHPO Ref.:   )

To : 	Project Director (CMHPO)

(Attn. Christine YEUNG)

### [by fax: 2127 4795 or email:cklyeung @healthbureau.gov.hk]

Your ref:

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

RFI (CE)

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply of Automated Endoscopic Reprocessor (hereinafter refers as the “System”) to The Chinese Medicine Hospital of Hong Kong (“CMHHK”) located at Pak Shing Kok in Tseung Kwan O.  The CMHPO therefore wishes to collect market information on thid System.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It is targeted to commence hospital services by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Automated Endoscopic Reprocessor that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Automated Endoscopic Reprocessor.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section G in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ________________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “Goods” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A. | Technical Requirements |  |  |
| 1. | Overall Requirements |  |  |
|  | The goods to be supplied is pass-through Automated Endoscopic Reprocessor (AER) [the “Goods”] comprising either following: (a) two (2) sets of AER with asynchronous operation of dual reprocessing basin; or (a) four (4) sets of AER with single reprocessing basin. |  |  |
|  | The Goods shall be able to reprocess flexible endoscope for cleaning and high-level disinfection purposes. |  |  |
|  | The Goods shall be delivered to and installed at Central Sterilization Service Unit (CSSU) on 3rd floor of The Chinese Medicine Hospital of Hong Kong (CMHHK) as designated by the CMHHK Operator and in compliance with the site requirements and conditions as specified in composite drawing at the Annex to these Technical Specifications. The Goods shall be installed at the specified wall opening(s) as indicated in the composite drawing. |  |  |
|  | General Features |  |  |
|  | The Goods shall be pass-through design to separate the clean zone and the dirty zone to prevent physical contamination across. |  |  |
|  | The Goods offered shall comply with the international standardisation of washer-disinfectors ISO 15883-1 and 15883-4 or equivalent standard. |  |  |
|  | For AER with asynchronous operation of dual reprocessing basin: |  |  |
|  | (a) 2 scopes processing in two different basins at the same time; and |  |  |
|  | (b) 4 bronchoscopes can be processing in the two basins at the same time with 4 scopes record and disinfection reports. |  |  |
|  | The Goods shall use pedal (hand-free) and / or control panel to open the lid. |  |  |
|  | The Goods shall have wireless bar code to capture the user information and endoscope recognition. |  |  |
|  | The Goods shall have LED screen in clean zone to display the scope information for individual basin spontaneously. |  |  |
|  | The Goods shall perform the self-disinfection cycle using programmable automatic start-up. |  |  |
|  | The Goods shall equip with air filtration system and filter system for the water filtered to 0.5 micron. |  |  |
|  | The Goods shall be able to trace the processes in both hardcopy format (using the integrated printer) and electronic format (using complete traceability management software). |  |  |
|  | The Goods shall use safe and validated single shot disinfectant chemical solutions, compatible with the various brands of endoscope available on the market (Olympus, Pentax, Fujifilm, Karl Storz, etc.). |  |  |
|  | The Goods shall have validated hook-ups connectors for all major brands of scopes, e.g. Olympus, Pentax and Fujifilm etc.. |  |  |
|  | The Goods shall provide Elevator Wire channel hook-up connector. |  |  |
|  | Each endoscope model has a unique hook-up block to ensure that proper disinfection can be achieved. |  |  |
|  | The Goods shall able to disinfect the soak-only devices with manufacturer certification, including not limited to following: |  |  |
|  | (a) Laryngoscope blades; |  |  |
|  | (b) Telescopes, probes; |  |  |
|  | (c) Cameras; |  |  |
|  | (d) Light cables; |  |  |
|  | (e) Transducers; |  |  |
|  | (f) Ultrasonic probes; |  |  |
|  | (g) Manometry probes; |  |  |
|  | (h) The video Laryngoscope; and |  |  |
|  | (i) Rigid Scope. |  |  |
|  | The Goods shall have continual individual channel monitoring: |  |  |
|  | (a) Monitor the flow rates in each channel; |  |  |
|  | (b) Monitor the blockage in each channel; and |  |  |
|  | (c) Monitor the connectivity in each channel. |  |  |
|  | The Goods shall perform leak test automatically at the cycle starting and check the correct pressure throughout the entire reprocessing cycle; and able to interrupt the cycle immediately if abnormality is found. |  |  |
|  | The Goods shall equip with dying cycle with automatic alcohol flush. |  |  |
|  | Tanks for the detergent/decontaminant and high-level sterilizing/disinfectant solutions are safe with no harmful emissions. |  |  |
|  | Design and Construct |  |  |
|  | Doors are interlocked and cannot be opened at the same time. |  |  |
|  | Door can only be opened on clean side upon completion of a full disinfection cycle. |  |  |
|  | Transparent doors to allow user to verify the progress, spray coverage and endoscope connections. |  |  |
|  | Emergency door opener to ensure scope can be removed from the basin in case of mechanical error of the electronic door during the reprocessing cycle. |  |  |
|  | All the maintenance and corrective services shall be conducted at the dirty zone to prevent the contamination of the clean zone. |  |  |
|  | The computer on the Goods is equipped with mirrored hard drives. Mirrored hard drives consist of a main hard drive and a second hard drive, which contains an exact copy of the first hard drive. This configuration provides a backup of the entire computer system. |  |  |
|  | Features related to Operation |  |  |
|  | Cassette shall be provided to place and remove endoscope from basin. The feature of cassette is included but not limited to following: |  |  |
|  | (a) Scope-specific port connector; |  |  |
|  | (c) Made from biocompatible polypropylene; |  |  |
|  | (d) Touch-free scope handling reduces cross-contamination; |  |  |
|  | (e) Provide protection of endoscopes during transport, the high of the tray shall be at least 1.5 inch. |  |  |
|  | Hook-up shall have a connector block with at least eight separate ports. |  |  |
|  | The hook-up connector block shall have labelled the channel name in order to prevent connection error. |  |  |
|  | All channels flush, disinfection and leak test shall be provided by one hook up connection. |  |  |
|  | At the end of every disinfection cycle, the system shall request the user to test the Minimum Recommended Concentration (MRC) of the disinfectant to ensure that the disinfection cycle efficiency. |  |  |
|  | The disinfectant sample shall be able to take from the built in sample port. |  |  |
|  | Upon completion of each reprocessing cycle, a report shall automatically prints out with not limited to following: |  |  |
|  | (a) Reprocessor serial number; |  |  |
|  | (b) Basin (for dual basin AER); |  |  |
|  | (c) Run number; |  |  |
|  | (d) Date and time; |  |  |
|  | (e) Hook-up used; |  |  |
|  | (f) Cycle result; |  |  |
|  | (g) Disinfectant temperature; |  |  |
|  | (h) Disinfectant contact time; |  |  |
|  | (i) Program run; |  |  |
|  | (j) Operator Load ID; |  |  |
|  | (k) Endoscope model / serial number ; |  |  |
|  | (l) Patient ID |  |  |
|  | (m) Process step times |  |  |
|  | (n) Errors |  |  |
|  | (o) Disinfectant MRC test results |  |  |
|  | Water Line Disinfection Cycle |  |  |
|  | There should be at least four water line disinfection programs available: |  |  |
|  | (a) Water line disinfection with 1-hour soak time, complete cycle is 2 hours; |  |  |
|  | (b) Water line disinfection with 4-hour soak time, complete cycle is 6 hour; |  |  |
|  | (c) Water line disinfection with 10-minute soak time, complete cycle is 1 hour; and |  |  |
|  | (d) Used to recover after failed water line disinfection or as a service test, complete cycle is 30 minutes. |  |  |
|  | Disinfection and Water Sampling Cycle |  |  |
|  | The reprocessing cycle shall be comprised of a number of specific steps designed to ensure proper endoscope disinfection. |  |  |
|  | High sensitivity leak test at the beginning of the cycle (before introduction of water). |  |  |
|  | A major leak (the diameter ≥0.01), the Goods shall alarm and the cycle shall be aborted before fluid ingress to avoid damage to endoscope. |  |  |
|  | Leak test checking throughout the cycle: |  |  |
|  | (a) Flushing; |  |  |
|  | (b) Washing; |  |  |
|  | (c) Rinsing; |  |  |
|  | (d) High level Disinfection; |  |  |
|  | (e) Automatic Alcohol Flush; and |  |  |
|  | (f) Air blowing with medical grade air. |  |  |
|  | The Goods shall allow for water sampling in basin during final rinse washing. |  |  |
|  | A sample port should be available inside the basin to ensure user can manipulate with collection of the water sample. |  |  |
|  | Software |  |  |
|  | The Goods shall equip with software program to be used for performing and running an endoscope cycle. |  |  |
|  | Program status shall be indicated on screen |  |  |
|  | The estimated percentage of the liquid remaining in the reservoirs shall be provided on the screen: |  |  |
|  | (a) Disinfectant(s); |  |  |
|  | (b) Alcohol; and |  |  |
|  | (c) Detergent. |  |  |
|  | All completed cycle logs can be viewed and reprinted at any time. |  |  |
|  | The software shall be able to manage the endoscopes, users and cycle information. It allows access to input data for report processing. |  |  |
|  | The software program with sort and search options to retrieve specific data from reprocessing cycle. |  |  |
|  | The software program allows setting up different user groups. |  |  |
|  | The data of chemistry change shall be recorded and stored: |  |  |
|  | (a) Type of chemistry change; |  |  |
|  | (b) Lot number of the chemistry changed; |  |  |
|  | (c) Time of chemistry change; |  |  |
|  | (d) Users who changed the chemistry; and |  |  |
|  | (e) The serial number of the reprocessor. |  |  |
|  | The software shall equip with delay start function. The endoscope can be placed in the basin, but the disinfection cycle will not start until pre-determined time input. |  |  |
|  | Chemistry and Water Requirements |  |  |
|  | The Goods shall use the detergent with well-documented efficiency for protein and biofilm removal. |  |  |
|  | The chemical sterilant or high-level disinfectant used for the Goods shall have manufacturer information such as safety data sheets and test reports supporting the efficacy claims against bacteria, fungi, viruses, bacterial spores, and mycobacteria. |  |  |
|  | Spillage kit for the disinfectant should be available to cope with accidental leakage. |  |  |
|  | The supplied waters: |  |  |
|  | (a) Potable water is minimum standard, Reversed Osmosis (RO) water is preferred but not mandatory; |  |  |
|  | (b) The incoming water to be filtered to 0.5 micron; |  |  |
|  | (c) Water pressure of 2 to 6 bar, with a minimum dynamic pressure of 2 bar at 5 L/min flow |  |  |
|  | (d) Minimum flow rate of 5 L/min |  |  |
| 9. | Air Supply Requirements |  |  |
| 9.1 | The suppled compressed air: |  |  |
|  | (a) Pressure: 4 and 10 bar; |  |  |
|  | (b) Minimum flow rate: 0.7 L/s; |  |  |
|  | (c) Maximum dew point is 7°C; |  |  |
|  | (d) Maximum oil concentration is 5 mg/m3; and |  |  |
|  | (e) Air usage is approximately 65 litres per cycle. |  |  |
| 10. | Equipment Safety |  |  |
| 10.1 | The Goods shall remain operational and within specifications throughout the voltage range of 220V ±6%, 50Hz ±2% single phase A.C. electrical supply. |  |  |
| 10.2 | The Goods offered shall comply with the safety requirements of IEC60601-1 or IEC61010-1 or equivalent standard. |  |  |
| 10.3 | The Goods offered shall comply with Electromagnetic Compatibility (EMC) Standards IEC 61326 or equivalent standard. |  |  |
|  | Implementation Services |  |  |
|  | The Supplier shall be responsible for connecting all waste outlets, data cable and electricity supply to the System and any modification works if required. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier shall note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
|  | The Supplier shall be responsible for the make good of ceiling, wall and floor in the installation site after completion of equipment installation. |  |  |
|  | Training |  |  |
|  | On-site maintenance and operational training shall be provided at no additional charges for maintenance and operation staff. |  |  |
|  | The Supplier shall be responsible to provide a session of on-site maintenance training to representatives of CMHHK upon request. The course shall cover at least basic theory of operation, circuit description, trouble-shooting technique, preventive maintenance procedures, calibration and alignment, adjustment. |  |  |
|  | The timetable and commencement dates for the training shall be advised at least one month prior to the commencement of the course. Detailed syllabuses shall be submitted for approval, upon request. The practical part of the training shall coincide with the installation and commissioning of the Goods. |  |  |
|  | Additional on-site application sessions shall also be arranged upon request by users during the warranty period without additional cost. |  |  |
|  | All training materials, such as course, video or slide presentation, diagrams and flowcharts, shall be available at the time of training. At least one copy of material, written in English, shall be provided to each trainee. |  |  |
|  | Documentation |  |  |
|  | All photocopies of operation and maintenance manuals shall be properly binded, stamped and certified as true copies of the original by the manufacturer. |  |  |
|  | At the time of delivery of the equipment, two Set(s) of the manufacturer’s original operation and maintenance manuals in English complete with principle of operations, operation instructions, trouble-shooting techniques, maintenance and calibration procedures, full parts list and full circuit diagrams levels shall be provided with the equipment ordered. |  |  |
|  | The maintenance manuals and checklist for preventive maintenance of the equipment shall be sent to the hospital after order placement but before the equipment delivery for documentation purpose. |  |  |
|  | The operation manual shall also be available as a standard softcopy PDF format. |  |  |
|  | Acceptance Tests |  |  |
|  | Safety Test For the purpose of this contract, the Goods shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Supplier with the witness of representative of CMHHK. The safety test will normally be conducted within 6 to 8 weeks after delivery and installation of the Goods. The date of completion by the Authority based upon the satisfactory result of such safety test. |  |  |
|  | If the Goods fail to comply with the safety standards listed in above clause, the Supplier shall investigate the causes of the failure and take appropriate corrective actions. The hospitals must be immediately informed of the failure. |  |  |
|  | In the case of medical equipment is reported with incident or safety problem, the Supplier shall:- - Attend to the call at site as soon as practically reasonable; - Conduct investigation to identify the cause of the incident causing the safety problem; - Advise the hospitals on the recommended actions to be taken based on investigation results; - Submit interim report and full investigation report to the hospitals within a time frame agreed by both parties. |  |  |
|  | Functional Test For the purpose of this Contract, the Goods shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. The Supplier shall provide all consumables for the evaluation. In the event that the equipment fails to conform requirements, the Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the Goods shall be determined by the Hospital based upon the satisfactory completion of such functional test. |  |  |
| 5. | The Supplier shall provide all testing instruments and consumables to conduct site acceptance tests. All testing instruments to be used for the acceptance test shall be calibrated and copies of calibration certificates or other supporting documents shall be forwarded to the representatives of CMHHK for records. |  |  |
|  | Indicative Warranty Service |  |  |
|  | The Supplier shall guarantee the equipment or any part thereof for a period of at least 12 months commencing from the date of acceptance of the equipment. The Supplier shall also replace faulty parts and provide both schedule and breakdown maintenance service by qualified maintenance personnel. In case of replacement of parts, they will be free of charge. |  |  |
|  | The Supplier shall submit as an essential part of the offer a yearly maintenance schedule during the warranty period indicating the number of preventive maintenance services required for ensuring a satisfactory performance of the equipment offered. Document, form, operation/service manual and/or manufacturer’s confirmation shall be submitted. If such information is not available, at least two times of preventive maintenance services shall be provided annually. The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours (09:00 – 18:00 hours Monday to Friday, excluding public holidays and 09:00 - 13:00 Saturday, excluding Public Holiday) |  |  |
|  | The Supplier shall be responsible to make good to the satisfaction of CMHHK Operator, any defects on the equipment due to improper workmanship, faulty design or component failure, which may arise within the warranty period of the equipment. |  |  |
|  | The Supplier shall guarantee the serviceable life span of the Goods for a period of not less than 7 years, from the date of acceptance of the Goods; otherwise, the Supplier shall guarantee to provide another set of at least similar (but not lower end) grade Goods to fulfil the 7 years’ service commitment. Supplier shall also guarantee to supply all spare parts for at least 10 years from that same date. |  |  |
|  | Upon notification by the CMHHK Operator of a defect in the operation of the equipment of part thereof, the Supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work being maintenance are not completed at the end of particular normal working period, subject to the CMHHK Operator’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the Supplier to carry on working until the particular maintenance task is completed. |  |  |
|  | Upon completion of the corrective maintenance works, the Supplier shall submit a report on the system breakdown investigation and corrective action taken. |  |  |
| 8. | Manufacturer-trained service engineers stayed locally to provide on-call services shall be available. |  |  |
| 9. | During warranty period, CM service call shall be unlimited and shall be carried out within 4 hours for emergency calls and no longer than 8 hours for other calls. |  |  |
| 10. | The Supplier shall assign designated person to support and answer the enquiry on the Goods’ performance and operation irregularities within 30 minutes. |  |  |
|  | Indicative Maintenance Service |  |  |
|  | The Supplier shall quote the charge for annual maintenance services after the warranty period within the serviceable life of the proposed Goods. |  |  |
|  | The Supplier shall submit a price list of all spare parts of the System chargeable to the CMHHK Operator. For spare parts not covered by the submitted prices, the Supplier must submit a quotation to the CMHHK Operator for consideration every time when spares are required. |  |  |
|  | The Supplier shall deploy properly trained service personnel to carry out the maintenance services and shall ensure that all necessary precautions for their safety are taken. |  |  |
|  | The Supplier shall provide free of additional charge corrective maintenance service for providing immediate repair service for the System and related equipment. |  |  |
|  | The maintenance services shall be carried out in accordance with the maintenance procedures as described in the relevant equipment services manuals. |  |  |
|  | Upon notification by the CMHHK Operator of a defect (departure from performance specifications) in the operation of the equipment of part thereof, the potential supplier shall perform the corrective maintenance within 48 hours upon request from the CMHHK Operator. This service shall include all necessary repairs, adjustment and replacement of parts to restore the equipment to its normal operational conditions in a time of no more than 3 working days. If such work is not completed at the end of particular normal working period, subject to the user’s agreement, the maintenance work will either be completed on next working day, or arrangement will be made for the supplier to carry on working until the particular maintenance task is completed. |  |  |
| 7. | Upon completion of the corrective maintenance works, the supplier shall submit a report on the equipment breakdown investigation result and corrective action taken. |  |  |
|  | Spare Parts |  |  |
|  | The supplier shall guarantee the availability of maintenance spare parts for the anticipated life of the Goods. Sufficient spare parts shall be held by the Supplier to cater for the maintenance during the warranty period. |  |  |
|  | The Supplier, in their tender submission, shall provide a comprehensive list of recommended spare parts with unit prices valid for at least one (1) year after expiry of warranty. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the Goods (if applicable) |  |  |
|  | #### Delivery of the Goods |  |  |
|  | #### Installation of the Goods |  |  |
|  | #### Implementation Goods (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the Goods has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Automated Endoscopic Reprocessor can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Automated Endoscopic Reprocessor does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Automated Endoscopic Reprocessor in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| ISO 15883-1 and 15883-4 or equivalent | The international standardisation of washer-disinfectors stated in clause 2.2. |  |  |  |
| IEC60601-1 or IEC61010-1 or equivalent | As per the requirements stated in clause 10.2 |  |  |  |
| IEC 61326 or equivalent | As per the requirements stated in clause 10.3 |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section H in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed Goods Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed Goods (leave blank if information is not available)

My/our proposed Goods was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed Goods (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the Goods, such as the requirements of ceiling mount support, power supply requirements, etc.)

END