P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Drying and Storage System for Endoscopes

to The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

(CMHPO Ref.  :       )

To : 	Project Director (CMHPO)

(Attn. Christine YEUNG)

### [by fax: 2127 4795 or email: cklyeung@healthbureau.gov.hk]

Your ref:

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

RFI (CE)

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and installation of Drying and Storage System for Endoscopes (hereinafter refers as the “System”) for the Chinese Medicine Hospital (“CMHHK”) located at Pak Shing Kok in Tseung Kwan O.  The CMHPO therefore wishes to collect market information on this System.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organization (“NPMO”) to operate the CMHHK. CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK. The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018. Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMHHK operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK. The CMHHK project has proceeded to the commissioning stage in 2021. The hospital services have been commenced by phases from 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Drying and Storage System for Endoscopes that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Drying and Storage System for Endoscopes.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section G in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ________________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “Goods” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
| A | Technical Requirements |  |  |
| 1 | Overall Requirements |  |  |
| 1.1 | Goods to be supplied: One (1) set of drying and storage system (the “System”) for endoscopes. |  |  |
| 2 | Drying and Storage System (“System”) |  |  |
| 2.1 | The System shall be designed for fast dry and storage of endoscopes with plasma technology to provide an active storage environment with aims to maintain the disinfected state of endoscope up to 31 days. The System is fully compatible with endoscopes from all major manufactures. |  |  |
| 2.2 | Quality Management System of the Manufacturer |  |  |
|  | (a) The manufacturer of the System shall have valid ISO13485 or equivalent. |  |  |
|  | (b) The System shall be certified with EN16442 or equivalent. |  |  |
| 2.3 | Marketing Authorization |  |  |
|  | The System shall have marketing authorization of at least one of the following region(s) and countr(ies) evidenced by the possession of: |  |  |
|  | (a) Relevant CE marking as evidences of its compliance with applicable legislation, technical and performance requirement of the European Union (“EU”) for selling in EU member states; or |  |  |
|  | (b) Relevant approval or clearance or exemption of the Food and Drug Administration (“FDA”) of the United States as evidences of its compliance with applicable legislation, technical and performance requirement of the United States for selling in the United States; or |  |  |
|  | (c) Appropriate national regulatory clearance from other members of the International Medical Device Regulators Forum (IMDRF). |  |  |
| 2.4 | General Description |  |  |
|  | The System shall include the following components (or equivalent): |  |  |
|  | (a) One set of drying and storage unit. (details refers A2.5) |  |  |
|  | (b) One set of endoscope connection set for connecting the endoscope to the drying and storage unit for drying endoscope channels. (details refers A2.6) |  |  |
|  | (c) One set of storage connection set for injecting plasma (containing ozone molecules) into the plasma bag during the storage cycle. (details refers A2.6) |  |  |
|  | (d) One pack of single-use plasma bags (at least 50 units/pack) to store disinfected endoscope and able to being injected with plasma containing ozone molecules during storage cycle. (details refers A2.7.1) |  |  |
|  | (e) One set of storage trolley with trays to place for endoscopes packed with plasma bag after storage cycle. (details refers A2.7.2) |  |  |
|  | (f) One pack of double-layer labels with 2 colors (at least 1,000 pcs/ pack) for sticking on the plasma bag to indicate the cleanliness status of the endoscope. (details refers A2.7.3) |  |  |
|  | (g) One set of label printer with SD memory card installed. (details refers A2.8) |  |  |
|  | (h) One set of barcode scanner for scanning the barcode of endoscope to identify the number and type of the endoscope and user ID. (details refers A2.9) |  |  |
|  | (i) One set of gas pressure regulator for medical air which is able to set the medical air inlet pressure to 3±0.2bar (43.51±2.9psi) from the medical air supply. |  |  |
|  | (j) One pack of shutters (at least 1,000 unit/pack) for sealing the orifice on the plasma bag at the end of the storage cycle. |  |  |
|  | (k) Endoscope Microbial Quality Test for 1 scope within a year. |  |  |
| 2.5 | Drying and Storage Unit |  |  |
| 2.5.1 | It includes not limited to following: |  |  |
|  | (a) Touch screen with control LED or equivalent: (i) size: at least 10cm; (ii) select and display steps of operation and alert message. |  |  |
|  | (b) One USB port at least for PC connection. |  |  |
|  | (c) One printer connector at least. |  |  |
|  | (d) One barcode scanner connector at least. |  |  |
|  | (e) One medical air inlet at least: medical air inlet at the rear of drying and storage unit: (i) minimal pressure: 3 bar (43.51 psi);and (ii) minimal flowrate: at least 60 L/min... |  |  |
|  | (g) Three gas/plasma outlets at least: at the front of drying and storage unit. |  |  |
| 2.5.2 | Program Requirements |  |  |
|  | (a) User identification: it should be performed by user login number. |  |  |
|  | (b) Endoscopes identification: it shall be performed by entering endoscope serial number attributing to endoscope type and shall be able to select the cycle. |  |  |
|  | (c) Cycle selection: (i) different drying cycles shall be able to select according to the types of endoscopes; (ii) shall able to select drying cycle, storage cycle or both drying and storage cycle. |  |  |
| 2.5.3 | Fast drying functions for endoscope channels of various types of scopes: |  |  |
|  | (a) Combination of different levels of flows (laminar and turbulent) and temperature to completely dry the endoscopes. |  |  |
|  | (b) Time of drying cycle within 3 minutes. |  |  |
|  | (c) Progress status of the drying cycle shall be shown on the screen. |  |  |
| 2.5.4 | It shall have not limited to following error message to alert user: (a) Incorrect pressure / air flow; (b) Absence of plasma; (c) Incorrect connection set plug; and (d) Overheating detected |  |  |
| 2.6 | Endoscope and Storage Connection Set (or equivalent) |  |  |
|  | (a) The tubing including plug shall be able for autoclave disinfection. |  |  |
|  | (b) It shall have different connection set to suit for different endoscope types of common brands. |  |  |
|  | (c) Using three different colors to identify the function of the connection sets: (i) for suction/operating channel; (ii) for air/water channel; and (iii) for water jet channel/storage. |  |  |
|  | (d) Storage connection set shall equip with pointed end of metallic cone-shaped adapter (or equivalent) to perforate the plasma bag during the storage cycle. |  |  |
| 2.7 | Storage Cycle (or equivalent) |  |  |
|  | The dried endoscope shall be put in the plasma bag and injected with plasma containing ozone molecules during storage cycle to make up an active storage environment to maintain the disinfected state of the endoscope. |  |  |
|  | (a) Up to 31 days. |  |  |
|  | (b) Time of storage cycle within 5 seconds. |  |  |
| 2.7.1 | Plasma Bag (or equivalent) |  |  |
|  | It shall be used to prevent infection risks upon storage and transportation. |  |  |
|  | (a) Single-use with adhesive tape at the opening of the bag for sealing purpose; |  |  |
|  | (b) Size: 60cm x 50cm (±10%); and |  |  |
|  | (c) Material: Low Density Polyethylene or equivalent. |  |  |
| 2.7.2 | Storage Trolley for Endoscopes |  |  |
|  | It shall be a medical-grade endoscope transportation trolley designed for safe hygienic, and secure transport of flexible endoscopes within hospital environments. |  |  |
|  | (a) It shall have four dual swivel casters with at least two are lockable. |  |  |
|  | (b) It shall be horizontal shelves design in order to prevent cross-contamination of each other. |  |  |
|  | (c) The trays shall be designed to slide into the horizontal shelves. |  |  |
|  | (d) It shall be able to store at least 8 flexible endoscopes which packed with plasma bag. |  |  |
|  | (e) It shall be constructed of high-quality stainless steel or other compatible materials to ensure durability, corrosion resistance, and ease of cleaning. |  |  |
|  | (f) All surfaces shall be smooth, non-porous, and easy to disinfect, with no sharp edges or joints that may harbor contaminants. |  |  |
|  | (g) It shall ensure safe handling of delicate endoscopes, preventing mechanical damage during movement. |  |  |
|  | (h) It shall be stable and not prone to tipping when fully loaded. |  |  |
|  | (i) It shall operate without excessive noise or vibration. |  |  |
| 2.7.3 | Double-layer Labels with 2 colors (or equivalent) |  |  |
|  | It indicates the cleanliness state of the endoscope: (a) One color such as green to indicate the endoscope stored in the plasma bag is under protected atmosphere and can be used. |  |  |
|  | (b) Another color such as red to indicate the endoscope is being used and soiled. User can put the soiled endoscope back to the plasma bag, refold the bag opening, and remove the green layer to reveal the red label to indicate the endoscope is soiled. |  |  |
| 2.8 | Label Printer with traceability labels (or equivalent) |  |  |
|  | The label printer to print out a label with traceable information. |  |  |
|  | (a) The label shall print out the information including below traceable information: |  |  |
|  | (i) date and time of the cycle start; |  |  |
|  | (ii) endoscope ID number with barcode; |  |  |
|  | (iii) user information; |  |  |
|  | (iv) completion of cycle information; and |  |  |
|  | (vi) expiry date to indicate the validity of storage. |  |  |
|  | (b) Print the label at the end of the storage cycle only or at the end of each cycle (both drying and storage). |  |  |
|  | (c) The printer shall be able to print at least 5 identical labels at the end of each cycle. |  |  |
|  | (d) The cycles history shall be saved as an Excel file (CSV type) on the memory card installed in the printer and it shall be accessed directly or via a local network. |  |  |
| 2.9 | Barcode Scanner (or equivalent) |  |  |
|  | (a) The barcode scanner shall be connected to the drying and storage unit via LEMO circular connectors. |  |  |
|  | (b) User can use the barcode scanner to scan the endoscope barcode for identification and user ID. |  |  |
| 2.10 | Electrical and Safety Requirement |  |  |
|  | (a) The System shall remain in full operation within the specification throughout the supply voltage range of 220V ± 6%, 50 Hz ± 2%, single phase A.C., and shall be fitted with, where applicable, a 13A power plug in compliance with BS 1363. |  |  |
|  | (b) The System shall comply with the safety requirements of IEC60601-1 / IEC61010-1 or equivalent. |  |  |
|  | (c) The System shall comply with electromagnetic compatibility (EMC) requirements of IEC60601-1-1 / IEC61326-1 or equivalent. |  |  |
|  | (d) The electrical installation, if any, provided by the successful Supplier shall be in compliance with the relevant requirements of the latest edition of “Code of Practice for the Electricity (Wiring) Regulations” issued by the Electrical and Mechanical Services Department (EMSD). |  |  |
| B | Implementation Services |  |  |
| 1 | The System shall be installed, tested and ready for use by the timeline specified in Part 4(k) with all costs from the date informed by the Hospital. |  |  |
| 2 | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
| 3 | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
| C | Training |  |  |
| 1 | Hands-on training for end users, including both clinical and non-clinical end users, shall be provided at no additional charges. |  |  |
| 2 | The timetable and commencement date for training shall be advised at least 3 months prior to the commencement of the course. Detailed syllabus is to be submitted at this time for approval. |  |  |
| 3 | Hardcopy and softcopy of the training material should be provided. |  |  |
| D | Documentation |  |  |
| 1 | The Supplier shall submit the documents listed as below in form of at least one hardcopy and one softcopy upon delivery of each equipment. All photocopies shall be properly blinded, stamped and certified as true copies of the original by the manufacturer. If the Supplier cannot provide softcopy of documents, the Supplier shall scan and provide the full documents as digital file (e.g. PDF file). The documents shall include, but not limited to, the latest version of (1) operation and maintenance (O&M) manual(s), in English or in Chinese complete with procedures of preventive maintenance (PM) and corrective maintenance (CM) and maintenance interval, (2) PM checklist(s) endorsed by equipment manufacturer (if PM is suggested by the manufacturer) and (3) installation manual/ method statement (if applicable) to user and hospital administration department. |  |  |
| 2 | Should any original equipment manufacturer products be included, the documents as specified in (a) above shall also be provided. |  |  |
| E | Acceptance Tests |  |  |
| 1 | Safety Test |  |  |
| 1.1 | The system shall be subject to a safety test after delivery and installation. Such test is to be carried out by the Authority Representative or his authorized nominee or nominees. The safety test will normally be conducted within 2 to 3 weeks after delivery and installation of the System. The date of completion of safety test for the System shall be determined by the Authority based upon the satisfactory result of such safety test. |  |  |
| 1.2 | The Hospital or its authorized nominee or nominees will not conduct safety tests until it has received full certification from the Supplier that the Supplier has conducted its own tests and found that the equipment meets the specifications in the Contract. |  |  |
| 2 | Functional Test |  |  |
| 2.1 | The system shall be subject to a functional test for its conformance with the operational and reliability requirements to the satisfaction of the user. In the event that the equipment fails to conform to the above stated requirements, the Supplier is required to carry out appropriate remedial measures and/or any rectification works, including replacement of the entire equipment, where deemed necessary. The date of acceptance of the System shall be determined by the Hospital based upon the satisfactory completion of such functional test. |  |  |
| 3 | The Supplier shall submit the acceptance test schedule, procedures, forms and testing method to the end-user for prior approval before the tests. |  |  |
| 4 | Acceptance test shall be performed on-site and the test report shall be submitted. All testing equipment shall be provided by Supplier and shall be calibrated. |  |  |
| F | Indicative Warranty Service |  |  |
| 1 | Warranty Maintenance |  |  |
| 1.1 | The System shall include a free warranty period of not less than 12 months starting from the date of Final Acceptance. |  |  |
| 1.2 | The Supplier shall provide a comprehensive, all-inclusive maintenance service, covering all spare parts, components, labor, and software support during the warranty period. |  |  |
| 1.3 | Upon notification of a fault, the Supplier shall provide on-site technical support within 4 hours (during office hours). |  |  |
| 2 | Preventive Maintenance (PM) |  |  |
| 2.1 | The Supplier shall perform routine Preventive Maintenance in accordance with the manufacturer’s recommendations (e.g., quarterly or half-yearly) to ensure the System operates at peak performance. |  |  |
| 2.2 | The preventive maintenance work shall be carried out as follows with no additional charge: Normal working hours 09:00 – 18:00 hours Monday to Friday, excluding public holidays 09:00 - 13:00 Saturday, excluding Public Holiday |  |  |
| 2.3 | All PM activities must follow a standardized checklist endorsed by the manufacturer and be submitted to the hospital administration for record-keeping. |  |  |
| 2.4 | PM services shall be scheduled in advance to minimize disruption to hospital clinical services. |  |  |
| 3 | Corrective Maintenance (CM) and Service Reports |  |  |
| 3.1 | The Supplier shall perform all necessary repairs, including the replacement of defective parts and software troubleshooting, to restore the System to full operational status. CM also includes the provision of software patches, version upgrades, and necessary toolkits/access codes for in-depth diagnosis. |  |  |
| 3.2 | The Supplier shall attend unlimited Breakdown Calls during the warranty period upon receipt of verbal or written notification from the CMHHK’s approved representative. Upon such notification, the Supplier shall, within the response time given, make such repairs or adjustments to, and replace such components/parts of, the System to restore the System back to its proper operating condition. |  |  |
| 3.3 | The service report shall be submitted within 24 hours from the time of incident/receipt of Breakdown Calls. |  |  |
| 3.4 | The Supplier shall provide the designated point of contact and make arrangements to enable maintenance representatives to receive requests timely for Corrective Maintenance service, when such need arises, twenty-four (24) hours a day, seven (7) days a week, including public holidays. |  |  |
| 3.5 | Service Reports After every maintenance visit (both PM and CM), the Supplier shall submit a detailed service report to the CMHHK. The report must include: - Date and time of the service call/request. - Description of the fault or maintenance task. - Details of parts replaced and labor hours spent. - Final status of the equipment and the signature of the attending technician and hospital witness. |  |  |

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Drying and Storage System for Endoscopes can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Drying and Storage System for Endoscopes  does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Drying and Storage System for Endoscopes  in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| ISO13485 or equivalent | Please refer to Section A2.2a for details. |  |  |  |
| EN16442 or equivalent | Please refer to Section A2.2b for details. |  |  |  |
| IEC 60601-1 or equivalent | Please refer to Section A2.10.b for details. |  |  |  |
| IEC 60601-1-2 or equivalent | Please refer to Section A2.10.c for details. |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 1 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

(b) Indicative Price Information for Selected Desirable Features (if applicable)

| Aspect | Description of Selected Desirable Features | Any Additional Charge to Total One-time Charge as Specified in Part 7(a) (Please tick whichever is applicable) |
| --- | --- | --- |
| Examples: |  |  |
| 1 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 1 hour < continuous operation ≤ 2 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 2 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: 2 hours < continuous operation ≤ 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |
| 3 | The System shall be capable of operating continuously for the below duration with a fully charged battery without the assistance of other power supply: continuous operation > 3 hours | □ No additional charge □ Require additional charge: HK$ _________ |

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section H in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END