P R O F O R M A

Request for Market Information (“RFI”) for 
Supply of Gait Analysis and Training System

for The Chinese Medicine Hospital of Hong Kong (“CMHHK”)

(CMHPO Ref.  :           )

To : 	Project Director (CMHPO)

(Attn. Rosana LAM)

### [by fax: 2127 4795 or email: rsllam@healthbureau.gov.hk]

Your ref: .: ( 1 ) in HHB/H/24/17/3/7/2/88

In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 9 of this Proforma.

Part 1 – Supplier’s Contact Details

From:

(Name of the Supplier): _______________________________________________________

(please fill in)

Name and Post of Contact person: _______________________________________________

(please fill in)

Email:___________________________       Telephone no.:________________________

(please fill in)					   (please fill in)

### -----------------------------------------------------------------------------------------------------------------

RFI (SE)

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.

Purpose and Background Information of the RFI

- Purpose

## Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the Supply of Gait Analysis and Training System (hereinafter refers as the “System”) for The Chinese Medicine Hospital of Hong Kong (“CMHHK”) located at Pak Shing Kok in Tseung Kwan O.  The CMHPO therefore wishes to collect market information on computed tomography systems.

- Background of the CMH Project

## The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK.  CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.

## The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services.

## To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMH operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It has been commenced hospital services by phases from December 2025.

## More information on the services provision and design of the CMHHK can be found in the following link:

## https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers

- If your company have more than one Gait Analysis and Training System that may meet the requirements of the System stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different Gait Analysis and Training System.

Part 2 – General Information of the System

| - Place of origin |  |
| --- | --- |
| - Name of manufacturer |  |
| - Address of the manufacturer’s factory or plant (“Manufacturing Plant”) |  |
| - Product name of the System |  |
| - Model number/ name/ version number of the System |  |
| - Authorised agent or distributor of the manufacturer in Hong Kong |  |
| - Packing (if applicable) |  |
| - Delivery method and route (where the place of origin is outside Hong Kong) |  |
| - Warranty period of the System (Please refer to section F in Part 3 for details of the warranty service requirements) | ____________ months from Acceptance of the System (Should not be less than 12 months) |
| - Expected serviceable life (Please specify any components of the System that cannot meet the serviceable life) | The System shall have a serviceable life of _______ years from its date of acceptance except the following components: ____________________________________ (Please also provide the expected life of these excluded components) |

Part 3 – Indicative Technical Requirements

Notes to Suppliers for Completion of Part 3

- Unless specified otherwise, the “System” in this Part 3 refers to section A1.1 below.

- The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.

- Please indicate, as a point by point compliance statement, whether your proposed System “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively.

- Where applicable, please quote the value of your proposed System in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement  (use additional sheet(s) if space is insufficient

- Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed computed tomography system against the corresponding indicative technical requirements.

| Column I | Column II | Column III | Column IV |
| --- | --- | --- | --- |
| Section | Technical Specification | Tick () the Appropriate Box (For aspects “Not Comply”, please also provide alternative proposal, if any) |  |
|  |  | Comply | Not Comply |
|  | General Requirements |  |  |
|  | System to be Supplied |  |  |
|  | The Gait Analysis and training System (“Gait Analysis System”) shall have the following components:- - a gait analysis and training treadmill with foot sensors, 3D camera, system software and integrated computer (“Training Treadmill”) as detailed in section 3 below; and - a dynamic Weight Relief System (“WR System”) that provides a dynamic weight support through the application of a sling to relieve the body weight of patient for walking training and fall prevention on the treadmill and helping to maintain patient in an upright posture as detailed in Section A3.13 below. |  |  |
|  | The Gait Analysis System shall be capable of performing gait analysis and gait training for walking and running, providing performance guidance through real-time feedback of gait parameters, offering virtual reality training, ensuring a safe walking and an upright posture with a dynamic weight relief system. |  |  |
|  | A total of three(3) sets of Gait Analysis System is required in this exercise. |  |  |
|  | Intended Use of the Gait Analysis System |  |  |
|  | The Gait Analysis System shall be equipped with load sensors in the tread belt for detection of patient’s movement and loading transfer during walking or running. |  |  |
|  | Functional and Performance Requirements of the Gait Analysis System |  |  |
|  | The Gait Analysis System shall be equipped with a 3D camera capable of detecting every single angular movement of the patient’s body, the position of the patient on the tread belt and the body joints of the patient without the need for attaching markers on the patient, and shall allow real-time reconstruction of the patient by digitizing it and transforming it into a virtual patient to be displayed on a digital monitor. |  |  |
|  | The Gait Analysis System shall have a large digital monitor of at least 48 inches at full HD resolution in front of the tread belt to function as a digital mirror to reflect the whole body image and real-time joint movement of the patient as a biofeedback display, and to provide a display of virtual reality environment for interactive and motivational training. |  |  |
|  | The Gait Analysis System shall have a speed control system activated by the click of a button for the tread belt speed, and shall be capable of accelerating and decelerating automatically by adapting to the motion action of the patient including the rhythm of steps and strides. |  |  |
|  | The Gait Analysis System shall be equipped with a computer with touchscreen control console for running software operation and for data entry and display. |  |  |
|  | The Gait Analysis System shall be capable of performing gait and run analysis, and gait training with real time evaluation and biofeedback of gait parameters through the interaction of load sensors in tread belt, 3D camera and software function. |  |  |
|  | The Gait Analysis System shall be capable of providing an interactive virtual reality environment and gamified task through display on the large digital monitor to enhance patient participation and motivation. |  |  |
|  | The Gait Analysis System shall have foot sensor applying to the back of the foot to measure gait parameters and ankle angles and shall synchronize wirelessly with the system software for data analysis and display. |  |  |
|  | Hardware Requirements of the Training Treadmill |  |  |
|  | The running surface of the Training Treadmill shall have a dimension of at least 160 cm (L) x 50 cm (W) to accommodate patients of different body types. |  |  |
|  | The Training Treadmill shall have a low access height of 15 cm or less. |  |  |
|  | The Training Treadmill shall have a speed range of 0 to at least 20 km/hour with an incremental step of 0.2 km/hour or lower. |  |  |
|  | The Training Treadmill shall have a reverse speed range of 0 to at least 5 km/hour with an incremental step of 0.2 km/hour or lower. |  |  |
|  | The Training Treadmill shall allow an elevation of 0% to at least 16%. |  |  |
|  | There shall be load detecting sensors under the running surface to detect load distribution. |  |  |
|  | The Training Treadmill shall support patient capacity of up to at least 150 kg. |  |  |
|  | The Training Treadmill shall have a dimension of 90 cm 5% (W) x 275 cm 5% (D) x 205 cm 5% (H). |  |  |
|  | 3D Camera Features of the Training Treadmill |  |  |
|  | The 3D camera shall be able to detect the patient in front of the tread belt automatically. |  |  |
|  | The 3D camera shall be equipped with two optics including a classic one for 2D shooting and one for infrared rays. |  |  |
|  | The 3D camera shall allow real-time reconstruction of the patient by digitalizing it and transforming it into a virtual image. |  |  |
|  | Foot Sensor Features of the Training Treadmill |  |  |
|  | The foot sensor shall be a small sensor applied on the back of the foot to measure gait parameters and ankle angles. |  |  |
|  | The foot sensor shall communicate with the treadmill through wireless connection. |  |  |
|  | The foot sensor shall be able to be used in stand-alone mode for over-ground walking in addition to treadmill walking. |  |  |
|  | The foot sensor shall be able to measure kinematic parameters of the foot in a dynamic situation during walking and running, and shall include dorsiflexion-plantarflexion and pronation-supination of the foot during heel strike and toe off. |  |  |
|  | Data measured by the foot sensor shall be able to synchronize automatically with the system software for data integration and display. |  |  |
|  | System Software Functions of the Training Treadmill |  |  |
|  | The system software shall have program setting function in terms of speed, elevation and cardio control. |  |  |
|  | The system software shall have gait and run analysis function with detection of length, speed and symmetries of the step, and range of motion of the trunk, hips and knees. |  |  |
|  | The system software shall have gait training function which shall allow therapist to reconstruct a mapping of patient’s pace through real time biofeedback correction of gait, and to set up targeted gait and postural parameters for gait correction and training. |  |  |
|  | The system software shall have test module function which shall offer at least 3 tests including walk/run test, 6 minutes walk test and Cooper treadmill test. |  |  |
|  | The system software shall have report function which shall allow the display of gait analysis parameters in terms of kinematic graphics and shall allow data export. |  |  |
|  | The system software shall have game module function which shall include virtual reality games for motivation and interactive walking and running training. |  |  |
|  | The system software shall have patient visualization function which shall be able to generate virtual simulation of patient through images captured by the 3D camera, and shall be able to provide the following features:- - The virtual simulation of patient shall be generated by the single 3D camera without the need for recording sideways and from the back; - The virtual simulated patient shall be able to synchronize with the real motion of the patient; - The virtual simulated patient shall be able to be viewed in 360 degrees and allow therapist to visualize patient’s gait in any planes of motion; and - Patient shall have the possibility of seeing his real face in frontal plane vision with the large display monitor functioning as a digital mirror. |  |  |
|  | The system software shall allow therapists to create, select and view patient data file and related information. |  |  |
|  | The system software shall be able to save the reports of assessment and training history in the patient data file. |  |  |
|  | Integrated Computer Requirements of the Training Treadmill |  |  |
|  | The central processor shall be i5 at 2.5GHz or higher. |  |  |
|  | The RAM shall be 4 GB or more. |  |  |
|  | The hard-disk shall have a capacity of 500 GB or higher. |  |  |
|  | The operating system shall be Windows 11 or higher. |  |  |
|  | Functional and Performance Requirements of the WR System |  |  |
|  | The WR System shall be a load support and relief system designed to work in synergy and fully integrate both ergonomically and electronically with the Training Treadmill. |  |  |
|  | The WR System shall provide dynamic support function for natural walking on Training Treadmill and fall prevention function with automatic treadmill stop. |  |  |
|  | The WR System shall make use of a hybrid relief system providing electronic and pneumatic control for the load support and relief. |  |  |
|  | The WR System shall have software interface to the control console of the Training Treadmill for regulation control of the amount of load support and relief in both static and dynamic way. |  |  |
|  | The WR System shall provide dynamic relief support with adjustable relief between 10 – 90 kg. |  |  |
|  | The WR System shall support patient capacity of at least 150 kg. |  |  |
|  | The WR System shall have dimensions of 120 cm 5% (W) x 230 cm 5% (D) x 265 cm 5% (H). |  |  |
|  | Safety and Product Standards |  |  |
|  | The Gait Analysis System shall comply with the latest edition of the safety standard of IEC60601-1 and IEC 60601-1-2, or their equivalent. |  |  |
|  | The Gait Analysis System shall comply with the latest edition of the safety standard of Directive 93/42/EEC of the European Union or equivalent. |  |  |
|  | The Gait Analysis System shall remain in full operation within the specification throughout the supply voltage range of 220V ±6%, 50Hz ±2% single phase A.C. |  |  |
|  | Single phase mains operated Gait Analysis System shall be fitted with a power plug suitable for the site condition. The plug shall comply with the standard of BS1363 for 13A plugs or equivalent. |  |  |
|  | The Gait Analysis System shall be effectively bonded to earth unless it is double insulated. The equipment shall be equipped with over-current protective cutout device. |  |  |
|  | All components of the Gait Analysis System shall be free of burrs, sharp edges, protrusions and other defects which may cause hazard to the patient or staff. All surfaces and edges shall be smooth and non-abrasive. |  |  |
|  | All motors, gears, chains, belts and flywheels of the Gait Analysis System, if equipped, shall be enclosed in protective covers and shall not be accessible by users during normal operation of the equipment. |  |  |
|  | The Training Treadmill shall be equipped with at least one push-button type emergency stop switch and at least one pull-cord type emergency stop switch, both of which shall be easily accessible by the users. |  |  |
|  | The requirements for flexible power cords of the Gait Analysis System including the marking requirements on the outer sheath, the requirements on color identification, conductors and insulation shall comply with the relevant standards. For example, the thickness of the insulation shall not be less than the requirements specified in IEC60227 or IEC60245 or EN 50525 or equivalent. |  |  |
|  | The Contractor shall comply with the provisions of the Factories and Industrial Undertakings (Lifting Appliances and Lifting Gear) Regulations (Chapter 59J of the Laws of Hong Kong) and provide a duly signed Form 4 “Certificate of Test and Thorough Examination of Lifting Appliances (Except Cranes, Crabs and Winches)” during equipment acceptance test stage; and Form 5 “Lifting Appliances Certificate of Results of Thorough Examinations in the Preceding Twelve Months” during Warranty Period for the WR System certified by a competent examiner1. |  |  |
|  | Implementation Services |  |  |
|  | The System shall be installed, tested and become ready for use by the timeline specified in Part 4(k) with all costs included. |  |  |
|  | Installation Requirements |  |  |
|  | Unless instructed otherwise by the Government/ CMHHK Operator, the three(3) sets of Gait Analysis System shall be installed at the following locations: Location; No. of set(s) to be installed. Satellite rehabilitation Room, 5/F; 1 set. Satellite rehabilitation Room, 6/F; 1 set. Satellite rehabilitation Room, 7/F; 1 set. |  |  |
|  | The Supplier shall be responsible to clear away all packing materials, demolished and unused structural materials to a legal place after delivery/installation of the equipment at no extra charges. |  |  |
|  | The Supplier should note that they will be held responsible for any damage to hospital property or that of the building contractor as may be caused during equipment transportation and installation. The Supplier should take all due measures to protect such property. |  |  |
|  | Training |  |  |
|  | Two (2) sessions of on-site operation and service training course on the operation and application of the Goods including without limitation full operation procedures of the Goods (including also the use of bundled software, programmes and applications, if any), routine maintenance and trouble-shooting techniques including where applicable basic theory of operation, circuit description, and techniques in setting up, calibration, alignment, trouble-shooting, fault diagnosis, preventive maintenance and corrective maintenance of the Goods, use of specialized tools and test instruments, etc. |  |  |
|  | Each training session shall be no less than two (2) hours and have a capacity to cater for at least ten (10) Trainees. |  |  |
|  | The above training course(s) shall be conducted on-site in CMHHK. hands on training should be carried out in parallel with the commissioning of the Goods. |  |  |
|  | Complete set of training materials for both operation and service training in hardcopy and softcopy should be provided. |  |  |
|  | Documentation |  |  |
|  | The Supplier shall provide the following manuals in hard copy before commencement of the acceptance tests: Two (2) sets of operation and maintenance manuals in English or in Traditional Chinese/Simplified Chinese (at least one set shall be original) shall include, but not limited to the following information under separate sections where applicable: • Description of the Goods; • Spare parts and special tools list - full parts list with details down to component level for maintenance purpose of the Goods; • Manufacturers’ certificates; • Safety precautions for operation and maintenance; • Operation instructions - principle of operations; • Maintenance instructions - trouble-shooting techniques; • Maintenance schedules; and • Drawing lists and drawings - block diagrams, circuit diagrams. |  |  |
|  | Should any original equipment manufacturer products be included, the documents as specified in (a) above shall also be provided. |  |  |
|  | Acceptance Tests |  |  |
|  | The acceptance tests shall include checking on materials used, safety device/features, structure strength, functional test and performance, and the test report shall be submitted. |  |  |
|  | It shall include step by step full functional tests to demonstrate that the functions and performance of each unit of the Goods and each part thereof are in full compliance with all requirements (such as functional, operational, reliability and safety requirements) of the Contract, including but not limited to the mandatory features and desirable features (to the extent such desirable features, if any, have been offered by the Contractor and accepted by the Government) set out in the Technical Specifications. |  |  |
|  | Where Gait Analysis System consist of several components, integration tests to ensure that all the components of each unit of the Gait Analysis System are interoperable and that each unit of the Goods is in compliance with the performance and functional requirements specified in the Technical Specifications. |  |  |
|  | Indicative Warranty Service |  |  |
|  | Warranty Maintenance |  |  |
|  | The supplier shall during the Warranty Period provide all of the following free warranty maintenance services at its own cost, including all parts and labour, to maintain the Goods in its full working conditions at no charge to the Government and the CMHHK Operator. |  |  |
|  | On-site preventive maintenance services at least one (1) time per annum. The on-site preventative maintenance services shall include without limitation (i) safety check, routine performance check and maintenance work; (ii) inspection and testing of the Goods (both functional and electrical safety tests shall be included); (iii) adjustment, calibration, cleaning and lubrication of the Goods; and (iv) replacement of defective parts of the Goods with brand new parts. |  |  |
|  | On-site corrective maintenance services shall include without limitation (i) inspection, electrical safety check, testing and diagnosis of faults; and (ii) carry out of such repair, adjustment, replacement of any unit of defective parts of the Goods to brand new parts; to restore the goods to proper working conditions. |  |  |
|  | Any replacement parts provided by the Contractor as well as the defective part(s) shall become the property of the CMHHK Operator. |  |  |
|  | The supplier shall provide all tools, instruments, software, spare parts and materials that are required for providing the free warranty maintenance services. |  |  |
|  | During the Warranty Period, the supplier undertakes and warrants that it shall attend to service calls within four (4) hours during normal business hours for the provision of onsite maintenance services. Supplier shall timely carry out such onsite maintenance services to restore the Goods to proper working conditions within two (2) working days from the receipt of the service call unless otherwise agreed by the CMHHK Operator. |  |  |
|  | The Contractor shall provide a hotline for fault reporting to ensure prompt fault attendance. |  |  |
|  | The free warranty maintenance services shall be carried out during normal business hours of a working day as follows: Monday to Friday : 09:00 to 17:00 Saturday : 09:00 to 13:00 |  |  |
|  | Indicative Maintenance Service |  |  |
|  | The Supplier shall quote the charge for annual maintenance services after the warranty period within the serviceable life of the proposed Goods. |  |  |
|  | The Supplier shall submit a price list of all spare parts of the Goods chargeable to the CMHHK Operator. For spare parts not covered by the submitted prices, the Supplier must submit a quotation to the CMHHK Operator for consideration every time when spares are required. |  |  |
|  | Spare parts |  |  |
|  | The Contractor shall guarantee the availability of maintenance spare parts for the Goods. Sufficient spare parts shall be held by the Contractor to cater for the maintenance during the Warranty Period. |  |  |
|  | The Contractor, in their tender submission, shall provide a comprehensive list of recommended spare parts with unit prices valid for at least one (1) year after expiry of warranty. |  |  |

Remark:

  - In accordance with the Factories and Industrial Undertakings (Lifting Appliances and Lifting Gear) Regulations (Chapter 59J of the Laws of Hong Kong), a“competent examiner”(合資格檢驗員) refers to an individual who is:

- appointed by the owner required by these regulations to ensure that the test and examination is carried out;

- a registered professional engineer registered under the Engineers Registration Ordinance (Cap.409 of the Laws of Hong Kong) within a relevant discipline specified by the Commissioner for Labour; and

- by reason of his qualifications, training and experience, competent to carry out the test and examination.

For full legislative context, please refer to https://www.elegislation.gov.hk/hk/cap59J

Part 4 – Implementation Plan

(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The System should be Ready for Use in the last month of the Implementation Plan.)

| Tasks of the Implementation Plan |  | Estimated Time Period for Performing the Tasks (The Order Date is set as Month 0) |  |
| --- | --- | --- | --- |
|  |  | Start (Month) | End (Month) |
|  | #### Order Date (i.e. the date of order placed by the Government, if any) | 0 | 0 |
|  | #### Submission of Site Preparation Information (if applicable) |  |  |
|  | #### Design of the System (if applicable) |  |  |
|  | #### Delivery of the System |  |  |
|  | #### Installation of the System |  |  |
|  | #### Implementation Services (Please refer to section B in Part 3 for details) |  |  |
|  | #### Delivery of Documentation (Please refer to section D in Part 3 for details) |  |  |
|  | #### Training (Please refer to section C in Part 3 for Details) |  |  |
|  | #### Acceptance Tests |  |  |
|  | #### Any other tasks considered necessary by your company (Please provide details, use separate sheet if space is insufficient): |  |  |
|  | #### System Ready for Use (i.e. the date when the System has passed all acceptance tests and accepted by the Government) | 0 |  |

Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)

(Note to Suppliers: Please indicate in the box below whether the proposed Gait Analysis and Training System can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed Gait Analysis and Training System does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed Gait Analysis and Training System in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

| Column I | Column II | Column III |  | Column IV |
| --- | --- | --- | --- | --- |
| International, National and other Recognised Standards or Certifications | Requirements | Comply with the Standard in Column I? |  | Comply with the following equivalent standard (If “No” in Column III) |
|  |  | Yes | No |  |
| IEC60601-1 or equivalent | As per the requirements stated in Section A4.1 |  |  |  |
| IEC 60601-1-2 or equivalent. | As per the requirements stated in Section A4.1 |  |  |  |
| Directive 93/42/EEC of the European Union or equivalent | As per the requirements stated in Section A4.2 |  |  |  |
| IEC60227 or IEC60245 or EN 50525 or equivalent | As per the requirements stated in Section A4.9 |  |  |  |
| Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify) |  |  |  |  |

Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)

(Note to Suppliers: Please advise whether your company and the proposed System have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.)

| Question | Licensing/Certification/Listing Information of the System | (Please tick in the appropriate box) |  |
| --- | --- | --- | --- |
|  |  | #Yes | No |
| ## 1 | ## Does your company have valid licence(s) to sell, deal with, possess and use irradiating apparatus in Hong Kong issued under the Radiation Ordinance (Chapter 303 of the Laws of Hong Kong) (“IA Licence”)? |  |  |
| 2 | Has the proposed System been listed in a valid IA Licence? |  |  |
| 3 | Dose the proposed System have marketing authorization of Food and Drug Administration (FDA) of the United States? |  |  |
| 4 | If the proposed System has marketing authorization of FDA, please specify below the type of marketing authorization (i.e. approval, clearance or exemption). _____________________ |  |  |
| 5 | Does the proposed System have marketing authorization of the European Union (EU) for affixing of CE marking on the product? |  |  |
| 6 | If the proposed System has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable). - *Declaration of conformity by the manufacturer; or - *Certificate of conformity issued by a notified body. |  |  |
| 7 | Does the proposed System have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”. Country / Region : _________________________________ |  |  |
| 8 | Has your proposed System been listed in the MDACS of the Department of Health? |  |  |
| 9 | What class of medical device is your proposed System (if applicable)? - EU : Class ______ - United States : Class ______ - Other country/region (please specify below): - Country/Region _________ - Class _________ |  |  |

#Please provide a copy of the licence/confirmation/certificate for reference.

## Part 7 – Indicative Price Information

(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

(a) Indicative Price Information for the System

| Item | Description | Estimated Quantity | Unit Price | Estimated Goods Price |
| --- | --- | --- | --- | --- |
|  |  |  | One-time Unit Price (HK$) | Estimated Goods Price for the Item specified opposite (HK$) |
|  |  | (a) | (b) | (c) = (a) x (b) |
| 1 | Supply, delivery, installation, testing and commissioning of the System and related accessories, as more particularly specified in section A1.1 in Part 3, including the provision of a minimum 12-months warranty period. | 3 set |  | (Please also provide breakdown cost for key components of the System, if any) |
| 2 | Provision of implementation services as detailed in section B in Part 3 | 1 lot |  |  |
| 3 | Provision of training services as detailed in section C in Part 3 | 2 courses |  |  |
| 4 | Documentation as detailed in section D in Part 3 | 1 lot |  |  |
| 5 | Other (please specify) | (please specify) |  |  |
| Total One-time Charge* (i.e. Sum of Estimated Goods Prices of Item 1- 5) |  |  |  |  |

Note: * The Total One-time Charge shall include one-year of warranty period.

## Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)

- Pursant to item 1 of Part 7(a) above, the proposed System shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.

- Indicative maintenance service requirements after the free warranty period are stipulated in section G in Part 3, which are subject to changes at the sole discretion of the Government

- It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the System may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.

- The annual maintenance charge within the serviceable life of the proposed System is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service.

- Indicative Maintenance Prices of the Proposed System

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| First 12-months period of maintenance service after the end of warranty period |  |
| Second 12-months period of maintenance service after the end of warranty period |  |
| Third 12-months period of maintenance service after the end of warranty period |  |
| Fourth 12-months period of maintenance service after the end of warranty period |  |

| Year | Annual Maintenance Charge (HK$ per annum) |
| --- | --- |
| Fifth 12-months period of maintenance service after the end of warranty period |  |
| Sixth 12-months period of maintenance service after the end of warranty period |  |
| Seventh 12-months period of maintenance service after the end of warranty period |  |

- Indicative Replacement Prices of System’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)

(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

| Item | Name of Items | Indicative Replacement Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |

- Indicative overtime charges for provision of maintenance services after office hours (if applicable)

(Office hours mean 9 am to 6 pm from Monday to Friday excluding public holidays)

| (a) | Rates of overtime charges for maintenance service outside the office hours | HK$ per hour |
| --- | --- | --- |
| (b) | Minimum service hour(s) per call | service hour(s) per call |

- Indicative Prices for Replacement of Other Spare Parts (if applicable)

| Item | Name of Items | Price (HK$/no.) | Indicative Replacement Frequency (e.g. once every 3 years) | Expected time for delivery (weeks) |
| --- | --- | --- | --- | --- |
| 1 |  |  |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |

- Indicative Price for Annual Support Services of Software (if applicable)

(Note to Suppliers: Please provide below annual charge for support services of the System’s software during the serviceable life of the System for the CMH Operator’s consideration.  The support services should include but not limited to:

- provision and renewal of software toolkits, access codes, passwords, software keys and hardware keys, etc. necessary for all kinds of adjustments, in-depth diagnosis and trouble shooting of the System; and

- version upgrade of the software.)

|  | (a) Free of charge during serviceable life |
| --- | --- |
|  | (b) Yearly cost at $__________________ |

Part 9 – Supplementary Information

- Number of proposed System Already Installed (leave blank if information is not available)

          - In Hong Kong :	 __________ sets

          - Globally		: __________ sets

- Year of Launch of the Proposed System (leave blank if information is not available)

My/our proposed System was first launched in the market in Year  _____________

- Pre-Installation Requirements of the Proposed System (if any)

(Pre-installation requirements may include any preparation work and provisions that are necessary for the installation of the System, such as the requirements of ceiling mount support, power supply requirements, etc.)

END