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P R O F O R M A
Request for Market Information (“RFI”) for 
Supply and Delivery of Hospital Bed and Accessories 
for The Chinese Medicine Hospital of Hong Kong(“CMHHK”)
(CMHPO Ref. : (1) in L/M to HHB/H/24/17/3/7/1/29)

To : 	Project Director (CMHPO)
[bookmark: _GoBack]	(Attn. Ms Wen CHAN, PO(CMHPO)2C)
[by fax: 2127 4795 or email: wspchan@healthbureau.gov.hk]
Your ref: (1) in L/M to HHB/H/24/17/3/7/1/29 
In response to the RFI of the CMHHK, my/our company, with contact details provided in Part 1 below, would like to provide the information and relevant supporting documents in Parts 2 to 10 of this Proforma.
Part 1 – Supplier’s Contact Details

From:  
(Name of the Supplier): _______________________________________________________
(please fill in)

Name and Post of Contact person: _______________________________________________
(please fill in)

Email:___________________________       Telephone no.:________________________
(please fill in)					   (please fill in)
-----------------------------------------------------------------------------------------------------------------

This document does not constitute any offer or invitation / solicitation of any offer in connection with the exercise described herein.  Neither this document nor any activities in connection therewith shall create any legal obligations or liabilities in any way on the part of the Health Bureau (HHB) or the Government of Hong Kong Special Administrative Region.  Neither this document nor anything contained herein shall form the basis of any contract or commitment whatsoever.  In responding to the RFI, a respondent shall be deemed to have agreed to all the terms of this Request for Market Information.


Purpose and Background Information of the RFI
1. Purpose
Chinese Medicine Hospital Project Office (“CMHPO”) of the Health Bureau (“HHB”) of the Government intends to invite a tender for the supply and delivery of Hospital Bed and Accessories (hereinafter refers as the “Goods”) for The Chinese Medicine Hospital of Hong Kong (“CMHHK”) located at Pak Shing Kok in Tseung Kwan O.  The CMHPO therefore wishes to collect market information on hospital bed and accessories. 

2. Background of the CMH Project
The Chief Executive announced in the 2014 Policy Address that the Government had decided to reserve a site in Tseung Kwan O for setting up a CMHHK. The 2017 Policy Address stated that the Government decided to finance the construction of the CMHHK and identify by way of tender a suitable non-profit-making organisation (“NPMO”) to operate the CMHHK. CMHHK will be owned by the Government and the selected NPMO will operate the CMHHK.  The CMHHK would be positioned as a flagship Chinese Medicine (“CM”) institution leading the development of CM services and Chinese medicines in Hong Kong. It will be a change driver, promoting service development, education and training, innovation and research, and facilitating collaboration with both local and international parties.
The CMHHK with provision of 400 beds will provide a comprehensive range of CM services. Service types include pure CM services, services with CM playing the predominant role in collaboration with Western Medicine (“WM”) and Integrated Chinese-Western Medicine (“ICWM”) services. The scope of service to be provided in the CMHHK covers inpatient, day-patient, outpatient and community outreach services. 
To take forward the planning and development of the project on CMHHK, a designated office i.e. CMHPO, was established under the Health Bureau (the former Food and Health Bureau) on 2 May 2018.  Hong Kong Baptist University (HKBU) was selected as the Contractor for the CMH operation. HKBU, as the Contractor, has incorporated a company limited by guarantee, namely HKBU Chinese Medicine Hospital Company Limited as the Operator to manage, operate and maintain the CMHHK.  The CMHHK project has proceeded to the commissioning stage in 2021. It has been commenced hospital services by phases from December 2025.
More information on the services provision and design of the CMHHK can be found in the following link:
https://www.healthbureau.gov.hk/en/press_and_publications/otherinfo/200900_cmhp/index.html

Note to Suppliers
1. If your company have more than one model of the hospital bed and accessories that may meet the requirements of the Goods stated in this Proforma, please complete and return, together with relevant supporting documents, one set of Proforma for each different model of the hospital bed and accessories.

Part 2 – General Information of the Goods
Item 1 – Electric Hospital Bed and Accessories
	[bookmark: _Hlk151107186]Item 1.1 - Electric Hospital Bed (“Electric Bed”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	



	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)




	Item 1.2 - General Mattress (“Mattress”) 

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)






	Item 1.3 - Lifting Pole (“Pole”) 

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)





	Item 1.4 - Intravenous Drip Pole (“IV Drip Pole”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods (Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)





	Item 1.5 - Oxygen Cylinder Carrier (“Carrier”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 _____________________________________
(Please also provide the expected life of these excluded components)





Item 2 – Paediatric Ward Bed and Accessories
	Item 2.1 – Paediatric Ward Bed (“Paed Bed”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	



	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)




	Item 2.2 - General Mattress (“Mattress”) 

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)






	Item 2.3 - Intravenous Drip Pole (“IV Drip Pole”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods (Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 ____________________________________
(Please also provide the expected life of these excluded components)





	Item 2.4 - Oxygen Cylinder Carrier (“Carrier”)

	(1) Place of origin
	

	(2) Name of manufacturer
	

	(3) Address of the manufacturer’s factory or plant (“Manufacturing Plant”)
	


	(4) Product name of the Goods
	

	(5) Model number/ name/ version number of the Goods
	

	(6) Authorised agent or distributor of the manufacturer in Hong Kong
	


	(7) Packing (if applicable)
	

	(8) Delivery method and route (where the place of origin is outside Hong Kong)
	


	(9) Warranty period of the Goods 
(Please refer to section F in Part 3 for details of the warranty service requirements)
	
____________ months from Acceptance of the Goods
(Should not be less than 12 months)

	(10) Expected serviceable life (Please specify any components of the Goods that cannot meet the serviceable life)
	The Goods shall have a serviceable life of _______ years from its date of acceptance except the following components:


 _____________________________________
(Please also provide the expected life of these excluded components)





Part 3 – Indicative Technical Requirements
Notes to Suppliers for Completion of Part 3
(a) Unless specified otherwise, the “Goods” in this Part 3 refers to Item 1 and Item 2 below.
(b) The indicative technical requirements are for the purpose of collecting market information only.  They are subject to changes and do not represent the final technical requirements of the intended tender.
(c) Please indicate, as a point by point compliance statement, whether your proposed Goods “Comply” or “Not Comply” with an indicative technical requirement stated in Column II by ticking () in the appropriate box under Column III and Column IV respectively. 
(d) Where applicable, please quote the value of your proposed Goods in either Column III (if “Comply”) or Column IV (if “Not Comply”) respectively against corresponding indicative technical requirement (use additional sheet(s) if space is insufficient)
(e) Please provide supporting documents (such as catalogues, user manual and/or operation manual, DICOM conformance statement, etc.) to illustrate the features of your proposed Hospital Bed and Accessories against the corresponding indicative technical requirements.


	Column
I
	Column 
II
	Column 
III
	Column
IV

	Section
	Technical Specification
	Tick () the Appropriate Box
(For aspects “Not Comply”, please also provide alternative proposal, if any)

	
	
	Comply
	Not Comply

	
	Technical Requirements

	
	Item 1 - Electric Hospital Bed and Accessories 

	(A)
	The Goods offered by the Contractor shall comply with the latest edition of the following standards, regulations and statutory requirements, where applicable:

	1.
	The Electric Hospital Bed shall have Conformité Européenne (“CE”) Mark. Documentary proof of compliance shall be submitted. 
	
	

	2.
	The Electric Hospital Bed shall comply with the technical requirements of IEC60601-1 and IEC60601-1-2 or equivalent standard. Additionally, it shall meet the basic safety and essential performance requirements specified in IEC60601-2-52 or equivalent standard. Documentary proof of compliance shall be submitted.
	
	

	3.
	Equipotential bonding shall be provided to the electric parts of the Bed in compliance with the relevant requirements of the latest edition of "Code of Practice for the Electricity (Wiring) Regulations", enforced by Electrical and Mechanical Services Department (“EMSD”).
	
	

	4. 
	The electrical parts of the Electric Hospital Bed shall be in compliance with the relevant safety requirements of the latest edition of the Electrical Products (Safety) Regulation under Electricity Ordinance, Cap. 406 and the latest edition of "Code of Practice for the Electricity (Wiring) Regulations", enforced by EMSD respectively.
	
	

	5. 
	The flexible cords fitted with the Electric Hospital Bed shall comply with the marking requirements on the outer sheath and the requirements on colour identification, conductors and insulation. For example, the thickness of the insulation shall not be less than the figures given in IEC60227 and IEC60245.
	
	

	(B)
	The Goods offered by the Contractor shall include the following:

	1.
	(a) Electric hospital beds (“Electric Bed”), referred to as Item 1.1 as detailed in Section B2 below;
(b) General mattresses (“Mattress”), referred to as Item 1.2  as detailed in Section B3 below;
(c) Lifting poles (“Pole”), referred to as Item 1.3 as detailed in Section B4 below;
(d) Intravenous drip poles (“IV Drip Pole”), referred to as Item 1.4 as detailed in Section B5 below; and 
(e) Oxygen cylinder carriers (“Carrier”), referred to as Item 1.5 as detailed in Section B6 below.
	
	

	2.
	 Item 1.1: Electric Hospital Bed (“Electric Bed”)
The Electric Bed shall include the following features:
	
	

	2.1.
	The safe working load shall be at least 200kg. 
	
	

	2.1.1
	The frame of the Electric Bed shall be made of durable and light weight metals with anti-rusting and anti-scratching coating.
	
	

	2.2
	The accessible parts and accessories of the Electric Bed shall be free of burrs, sharp edges, protrusions and other defects which may cause hazard to patients and operators.
	
	

	2.3
	The Electric Bed shall be operated in 220V +/- 6%, 50Hz +/- 2%, 1-phase AC electricity supply with BS1363 13A plug fitted.
	
	

	2.4
	The Electric Bed shall be electrically operated and powered by mains supply, with an emergency backup battery. The operation of the Bed by mains supply shall be independent of the backup battery and vice versa.
	
	

	2.5
	The backup battery shall be rechargeable of solid type or sealed type.
	
	

	2.6
	The charger for the backup battery shall be equipped with an automatic cut off function to protect battery over-charging.
	
	

	2.7
	The Electric Bed shall be capable of operating for at least three (3) hours when powered by backup battery in standby mode.
	
	

	2.8
	The electrical parts of the Electric Bed such as control circuit and motor shall be operated by extra low voltage.
	
	

	2.9
	The electrical parts of the Electric Bed shall have ingress protection of at least IPX4.
	
	

	2.10
	The Electric Bed shall be equipped with an over-current protective cutout device.
	
	

	2.11
	The Electric Bed shall be effectively bonded to earth unless it is double insulated.
	
	

	2.12
	All motors and gears of the Electric Bed shall be enclosed in protective covers and shall not be accessible by users or patients during normal operation of the Bed.
	
	

	2.13
	The height and inclination the Electric Bed shall remain in their original positions when power to the equipment is terminated.
	
	

	2.14
	Dimension and Weight 
	
	

	2.14.1
	The overall length of the Electric Bed (excluding extension) shall not exceed 2300mm.
	
	

	2.14.2
	The overall weight of the electric bed shall not be heavier than 175 kg.  
	
	

	2.14.3
	The patient surface dimensions shall be at least 1900mm x 800mm.
	
	

	2.14.4
	The Electric Bed shall be designed with an extendable feature that includes a solid platform, allowing for an extension of at least 100 mm. This extension shall be equipped with a locking mechanism to ensure stability and safety during use.
	
	

	2.14.5
	The overall width of the Electric Bed including side rails but excluding revolving protective bumpers shall be within the range of 950mm to 1100mm.
	
	

	2.14.6
	The height movement of the Electric Bed measured from the floor to the top of the mattress support platform shall have the following measurement:
	
	

	2.14.6.1
	The highest position of the Electric Bed measured from the floor to the top of the mattress support platform, without mattress shall be in the range of 750mm to 900mm.
	
	

	2.14.6.2
	The lowest position of the Electric Bed measured from the floor to the top of the mattress support platform, without mattress shall be less than 450mm.
	
	

	2.14.7
	The under-bed space clearance of the electric bed frame shall be not less than 120mm height measured from the floor to the lowest surface of the bed frame.
	
	

	2.15
	Power Movements
	
	

	215.1
	The Trendelenburg and reverse Trendelenburg movement shall each allow for a minimum angle of 12 degrees.
	
	

	2.15.2
	The backrest of the Electric Bed shall allow for continuous upward adjustment with a minimum inclination of +60 degree and be capable of being stopped at any desired angle.
	
	

	2.16
	Construction
Each Electric Bed shall include the following components:
(a) Mattress support platform as detailed in Section B2.16.1 below;
(b) Headboard and footboard as detailed in Section B2.16.2 below;
(c) Side rails as detailed in Section B2.16.3 below;
(d) Bumpers as detailed in Section B2.16.4 below;
(e) Receptacles as detailed in Section B2.16.5 below;
(f) CPR Lever as detailed in Section B2.16.6 below;
(g) Castors and brakes as detailed in Section B2.16.7 below; and
(h) Control as detailed in Section B2.16.8 below.
	
	

	2.16.1
	Mattress Support Platform
	
	

	2.16.1.1
	The mattress support platform shall be flat (except minor lateral curvature for mattress retaining) and solid with smooth-finishing. Web like or grated support platform will not be considered.
	
	

	2.16.1.2
	The mattress shall be fire retardant and meet the safety requirement of BS 7177:2008+Al: 2011 (Medium Hazard) when tested in accordance with BS EN 597-1:2015 (Ignition source 0), BS EN 597-2:2015 (Ignition source 1) and BS 6807:2006 Section 2 (Ignition source 5), or equivalent. Documentary proof of compliance shall be submitted. 
	
	

	2.16.2
	Headboard and Footboard
	
	

	2.16.2.1
	The headboard and footboard shall be rectangular-like, rigid and easy-to-clean.
	
	

	2.16.2.2
	The headboard and footboard shall be removable as single, complete units from the bed frame, respectively, without using tools.
	
	

	2.16.2.3
	The headboard and footboard shall be easily attached and detached from the bed frame to facilitate fast and smooth cardiopulmonary resuscitation (“CPR”).
	
	

	2.16.2.4
	When footboard is mounted to the Electric Bed frame, the longitudinal stability shall comply with IEC60601-2-52 or equivalent standard.
	
	

	2.16.3
	Side Rails
	
	

	2.16.3.1
	The Electric Bed shall provide one (1) pair of two-piece segmented type side rails at each side of the Bed, resulting in a total of two (2) pairs of side rails.
	
	

	2.16.3.2
	The side rails shall be made of acrylonitrile butadiene styrene plastic or functionally equivalent material.
	
	

	2.16.3.3
	The height of the top edge of the side rails shall be at least 350mm when in the raised position above the mattress support platform.
	
	

	2.16.3.4
	When the side rails are raised to the locked position, height of the top edge of the side rails shall be at least 250mm above the mattress without compression.
	
	

	2.16.3.5
	The side rails of the Electric Bed shall allow hanging of drainage bag(s) on each side.
	
	

	2.16.3.6
	The Electric bed shall have the following fall-prevention releasing mechanism on the side rails:
(a) The side rails shall be designed to prevent patients from unlocking them while leaning against the rails;
(b) The side rails shall be soft-drop type to prevent 
entrapments; 
(c) When the side rail is lowered down, it shall remain     
securely attached and not crashed onto the bumpers; and
(d) The releasing mechanism of the side rails shall not require the operator to exert excessive force in awkward movement such as exerting force in two different directions on pressing the releasing button by fingers.
	
	

	2.16.3.7
	The side rails shall be designed in accordance with IEC 60601-2-52 or equivalent standard, to prevent from trapping the head, body and limbs of the operator, patient or carer.
	
	

	2.16.4
	Bumpers
	
	

	2.16.4.1
	The Electric Bed shall provide four (4) revolving protective bumpers which shall be stain-resistant at the four (4) corners of the Bed.
	
	

	2.16.5
	Receptacles
	
	

	2.6.5.1
	There shall be at least four (4) fixed receptacles for intravenous (“IV”) drip pole near the four (4) corners of the Electric Bed, and two (2) fixed receptacles for lifting pole near the head end corners. Each receptacle shall feature an anti-rotation slot for stability.
	
	

	2.16.6
	CPR Lever
	
	

	2.16.6.1
	A manual CPR lever or functionally equivalent design shall be provided. The CPR lever activation shall require no excessive force from the operator.
	
	

	2.16.7
	Castors and Brakes 
	
	

	2.16.7.1
	The Electric Bed shall have four (4) swivelling twin-wheel castors with diameter not less than 125mm. At least one of the castors shall have steering capability.
	
	

	2.16.7.2
	All castors of the Electric Bed can be locked simultaneously by a central braking pedal or equivalent device.
	
	

	2.16.7.3
	The quality of castors of the Electric Bed shall be in compliance with EN12531: 1999 or equivalent international, national and other recognised standards or certifications.
	
	

	2.16.8
	Controls 
	
	

	2.16.8.1
	The positioning and functioning of the Electric Bed shall be controlled by:
(a)	at least one (1) wired integrated control panel at the foot end of the Electric Bed;
(b)	at least one (1) integrated control panel on each outer side of the head side rails of the Electric Bed; and
(c)	at least one (1) integrated patient control (“IPC”) on each inner part of the head side rails.
	
	

	2.16.8.2
	The integrated control panel shall consist of the following control elements or functionally equivalent control:
(a) (a) Trendelenburg tilt button;
(b) (b) Reverse Trendelenburg tilt button;
(c) (c)Bed height adjustment lock indicator (ON-locked);
(d) (d) Bed height adjustment button (up/down);
(e) (e)Backrest adjustment lock indicator (ON- locked);
(f) (f) Backrest adjustment button (up/down);
(g) (g) Thigh section adjustment lock indicator (ON-locked);
(h) (h) Thigh section adjustment button (up/down);
(i) (i) Backup battery charging status indicator;
(j) (j) Emergency CPR position button.
	
	

	2.16.8.2.1
	The wired integrated control panel shall be equipped with a deep U-shaped or deep ear-shaper arch or functionally equivalent design with no sharp edges to fit onto the contour of the footboard.
	
	

	2.16.8.3
	The IPC shall consist of the following control elements to facilitate the positioning of the Electric Bed by patients:
(a)	Thighrest adjustment buttons (up/down); and
(b)	Backrest adjustment button (up/down).
	
	

	2.16.8.4
	The IPC shall not be capable of controlling the bed-tilting functions for Trendelenburg position, reverse Trendelenburg position and chair position while in the Trendelenburg and reverse Trendelenburg positions.
	
	

	2.16.8.5
	There shall be a locking mechanism for the controls of the Electric Bed to prevent unauthorized or unintentional activation of electric bed movements.
	
	

	2.16.8.6
	The controls of the Electric Bed shall have automatic protection and switch off the actuators when either they reach the limits of travel, or are overloaded. Functionally equivalent design is acceptable.
	
	

	3.
	Item 1.2: General Mattress (“Mattress”)
	
	

	3.1
	Weight and Dimension
	
	

	3.1.1
	The weight of the Mattress shall not exceed 10kg, and it must be compatible with the Electric Bed to ensure proper fit and functionality.
	
	

	3.1.2
	The mattress shall fit the mattress support platform to avoid the patient entrapment hazard between the side rails and the Mattress; and the headboard and footboard with the Mattress end.
	
	

	3.1.3
	The thickness of the Mattress shall be at least 100mm.
	
	

	3.2
	Construction
	
	

	3.2.1
	It shall be foam-filled completed with a mattress cover. 
	
	

	3.2.2
	The Mattress core shall be made from single-layer flexible, combustion modified Polyurethane (“PUR”) foam or equivalent.
	
	

	3.2.3
	The Mattress cover shall be woven / knitted nylon with PUR waterproof coating.
	
	

	3.2.4
	The coating of Mattress cover shall not be affected by common cleaning materials such as hospital disinfectants, lubricants, alcohol, soap and water that may be in contact or used for cleaning.
	
	

	3.2.5
	The Mattress cover shall be wrinkle free.
	
	

	4.
	Item 1.3: Lifting Pole (“Pole”)
	
	

	4.1
	The Pole shall be made of powder coated steel or equivalent material with a plastic triangular hand grip and adjustable strap.
	
	

	4.2
	The Pole shall be capable of fitting into the Pole receptacle located at head end of the Electric Bed without using tool and without movement.
	
	

	4.3
	A non-staining label shall be affixed onto the lifting pole indicating its weight and safe working load.
	
	

	4.4
	The design and size of the strap loop on the lifting pole handle shall be compatible with the Pole rod. The strap loop shall not be able to be removed from the rod through manual manipulation.
	
	

	4.5
	A position marking or equivalent mechanism shall be available at the foot end of the Pole to indicate that the Pole is fully and securely inserted into the receptacle.
	
	

	4.6
	The Pole shall be secure in place with no loose movement when supporting patient to a sitting position. 
	
	

	4.7
	The anchoring belt of the hand grip shall be able to firmly fasten at the desired length and fixed at various positions on the top portion of the Pole after adjustment.
	
	

	5.
	Item 1.4: Intravenous Drip Pole (“IV Drip Pole”)
	
	

	5.1
	A one-hand operated telescopic IV Drip Pole made of anti-corrosive materials with at least two (2) open mouth U-shaped or C-shaped hooks.
	
	

	5.2
	The safe working load of the hook shall be at least 1 kg each. 
	
	

	5.3
	The height of the IV Drip Pole shall be telescopically adjustable and shall be fixed at the desired height when the operation lever is released. Functionally equivalent design is also acceptable.
	
	

	5.4
	The telescopic movement shall be adjustable without excessive force and trapping the operation’s fingers. The handheld component for telescopic adjustment shall be non-slippery for firm grasping.
	
	

	5.5
	The IV Drip Pole shall fit securely into the IV drip receptacles that are welded (or locked) at the head and foot corners of the Electric Bed without loose movement and without using any tools.
	
	

	5.6
	The IV Drip Pole shall be located without any obstacle to the lateral side of the Electric Bed to ensure clear pathway for patient transfer.
	
	

	5.7
	The IV Drip Pole shall be located without tilting forward during powered movement of the Electric Bed head section.
	
	

	6.
	Item 1.5: Oxygen Cylinder Carrier (“Carrier”)
	
	

	6.1
	The Carrier is made of anti-corrosive materials with vertical design that shall be fitted with the IV Drip Pole receptacles.
	
	

	6.2 
	Each Carrier shall be capable of accommodating an oxygen cylinder of at least 140mm in diameter putting vertically.
	
	

	6.3
	The Carrier shall be secure in place with no loose movement when oxygen cylinder is placed.
	
	

	(C)
	Training
The Contractor shall provide Training meeting the following requirements:
	
	

	1.
	At least two (2) on-site comprehensive equipment operation and maintenance training courses shall be provided to the CMHHK Operator’s operational personnel and maintenance staff within two (2) weeks after completion of Acceptance Test.
	
	

	2.
	The training courses shall enable CMHHK Operator’s operational personnel and maintenance staff acquires knowledge on day-to-day operation, safety knowledge, routine maintenance and fault diagnosis.
	
	

	3.
	The training shall be conducted in Cantonese with English terms.  All training and training materials provided shall be in Traditional Chinese or English
	
	

	(D)
	Documentation
The Contractor shall provide operation and maintenance (“O&M”) manuals including requirements as below
	
	

	1.
	The Contractor shall submit at least two (2) sets of the manufacturer’s O&M manuals in English or in Traditional Chinese/Simplified Chinese as directed by the Government within two (2) weeks after completion of Acceptance Test. The Contractor shall submit the documentation in form of CD/DVD.
	
	

	2.
	The content of the O&M manuals shall include, but not limited to the following information under separate sections where applicable:
(a) (a) description of the Goods
(b) (b) spare parts and consumables, and special tools list
(c) (c) manufacturers’ certificates
(d) (d) safety precautions for operation and maintenance
(e) (e)operation instructions
(f) (f) maintenance instructions
(g) (g) maintenance schedules; and
(h) (h) drawing lists and drawings
	
	

	(E)
	Acceptance Tests
	
	

	1.
	On the date(s) specified in the Acceptance Test Plan, the Contractor shall demonstrate, in the presence of the Government Representative, that the Goods is properly tested by the Contractor; in compliance with all mandatory features as stated in the Technical Specifications.  
	
	

	2.
	The Contractor shall conduct a step-by-step testing of each mandatory features as stated in the Technical Specifications. The Government Representative may require the Contractor to make any amendments or adjustments during the tests or request to conduct further reasonable tests.  The Contractor shall be responsible for the provision of all the test equipment and test gear, without generated any additional cost to the Government and the CMHHK Operator.
	
	

	3.
	All tests shall be witnessed by the Government Representative and a minimum of seven (7) days written notice shall be given to the Government Representative before commencement of testing.
	
	

	4.
	The Contractor shall supply, at his own expense, all other necessary materials, as well as other necessary equipment and materials sufficient to complete the testing including repeated tests to the satisfaction of the Government Representative.  
	
	

	5.
	Unless otherwise indicated, the Contractor shall include the cost of all tests, necessary instrument, plant, supervision and labour both at works and on site. All instrument must be properly calibrated, and the accuracy of the test instrument shall be demonstrated where so directed by the Government Representative.  The Contractor shall also carry out all such additional tests as the Government Representative may require to establish that the Goods meets the requirements of the Technical Specifications.
	
	

	6.
	Any defects of workmanship, materials and performance, mal-adjustments or other irregularities which become apparent during the tests shall be rectified by the Contractor at his own expense and the tests shall be repeated at the Contractor's expense to the satisfaction and acceptance of the Government Representative.
	
	

	7.
	All tests, including those to comply with statutory requirements, shall be properly certified in a manner to be agreed with the Government Representative and triplicate copies of all certificates shall be submitted to the Government Representative and appropriate authorities on completion.
	
	

	8.
	The Goods shall be subject to functional and safety testing for their conformance with the operational and reliability requirements.  The tests shall comply with requirements of IEC 60601-1, IEC 60601-1-2 and IEC60601-2-52 or equivalent.
	
	

	9.
	Only when the Goods has passed all the tests and so certified and all test figures and other relevant information have been recorded in the prescribed manner and accepted by the Government Representative shall the Goods be considered fit for handing over to the Government Representative and the CMHHK Operator.
	
	

	(F)
	Indicative Warranty Service
	
	

	[bookmark: _Hlk151735195]1
	The Contractor shall provide information to the Government in writing within 24 hours on (a) faults and/or deficiencies known to the Contractor found in Goods supplied by the Contractor which are similar to the Government’s Goods and (b) methods developed to deal with such faults, and/or deficiencies.
	
	

	2
	Where the same fault or deficiency occurs more than one time within three (3) months, the Contractor shall provide assistance to the Government and the CMHHK Operator in finding the cause(s) and the remedies of such fault or deficiency and preventing further occurrences.
	
	

	3.
	The Contractor shall conduct safety investigation and provide investigation report with safety recommendations when there is a safety incident happened within the Warranty Period of the Goods.
	
	

	4.
	Preventive Maintenance
	
	

	4.1
	The Contractor shall provide preventive maintenance services according to manufacturer’s instructions / recommendations, or at least one (1) time per year if not specified, within the Warranty Period, covering full inspection, lubrication and testing in order to keep the Goods in good operating condition and in accordance with the Technical Specifications. The scope of maintenance shall include the following activities: 

(i)	checking and performing safety tests for ensuring satisfactory performance of the Goods;
(ii)	checking and performing all necessary repairs and adjustments of the Goods; and
(iii)	lubricating all components and replacement as necessary, to maintain the Goods in good order including supply of material and labour.
	
	

	4.2
	One (1) month before the start of the Warranty Period, the Contractor shall submit for the Government’s approval a preventive maintenance services schedule for the Goods.  The scheduled services should normally be carried out during normal business hours where other operation of the CMHHK shall not be interfered. The Contractor shall follow the schedule approved by the Government.
	
	

	4.3
	The Contractor shall perform an inspection to check over the Goods and submit a report to the Government within one (1) month before the end of the Warranty Period. The report shall include present condition of the Goods, any irregularities or faults found with cause of the fault identified, remedial actions taken or to be taken, parts repaired/replaced, and any follow-up actions or recommendations.
	
	

	4.4
	In the event of any irregularities identified during the preventive maintenance, the Contractor shall take remedial actions and report to the Government and the CMHHK Operator immediately.  The Contractor shall, based on the findings from the preventive maintenance, draw the attention of the Government and the CMHHK Operator immediately to take any preventive measures or appropriate action at an early stage in order to avoid major faults or degradation in performance of the Goods.
	
	

	5.
	Corrective Maintenance
	
	

	5.1
	The Contractor shall provide a hotline for fault reporting to ensure prompt fault attendance. The faults and complaints response time shall be within two (2) working days from the reporting of fault to the Contractor.
	
	

	5.2
	The Contractor shall rectify faults and perform replacement of parts, adjustments and calibrations as may be necessary to ensure the Goods works properly in accordance with the Technical Specifications.
	
	

	5.3
	In the event that the Contractor decides to replace a component of the Goods, the replaced component shall become the property of the CMHHK Operator.
	
	

	6.
	In case of breakdown of the Goods in the course of operation, or the maintenance of the Goods by the Contractor within the Warranty Period, all associated works, and the possible implications shall be borne by the Contractor. 
	
	

	7.
	During the Warranty Period, the maintenance staff of the Contractor shall complete the site record “Maintenance Log Book” in either English or Chinese after each on-site visit.  The entries shall give a full report of the works undertaken during the attendance, including description of fault, zone/areas affected by the fault, cause of fault, remedial actions taken or to be taken, parts repaired/replaced, any follow-up actions or recommendations.
	
	

	8.
	The Contractor shall inform the Government and the CMHHK Operator immediately of any intention by it or the relevant manufacturers to cease manufacture of any Spare Parts prior to any such cessation.  The Contractor shall, at its own cost, be responsible to keep sufficient stock of Spare Parts, and shall ensure that they are fully functional and in good working conditions.

	
	

	
	Item 2 - Paediatric Ward Bed and Accessories 

	(A)
	The Goods offered by the Contractor shall comply with the latest edition of the following standards, regulations and statutory requirements, where applicable: 

	1.
	The Paediatric Ward Bed shall comply with the technical requirements of IEC60601-2-52 or equivalent standard.  Documentary poof of compliance shall be submitted. 
	
	

	(B)
	The Goods offered by the Contractor shall include the following:
	

	1.
	(a) Paediatric Ward Bed (“Paed Bed”), referred to as Item 2.1 as detailed in Section B2 below;
(b) General mattresses (“Mattress”), referred to as Item 2.2  as detailed in Section B3 below;
(c) Intravenous drip poles (“IV Drip Pole”), referred to as Item 2.3 as detailed in Section B4 below; and 

(d) Oxygen cylinder carriers (“Carrier”), referred to as Item 2.4 as detailed in Section B5 below.
	
	

	2.
	Item 2.1: Paediatric Ward Bed (“Paed Bed”) 
The Paed Bed offered by the Contractor shall include the following features:
	
	

	2.1
	It shall be designed for hospital use, to assist in monitoring infants and young children.
	
	

	2.2
	It shall be manually operated with ergonomic crank handles for adjustment.
	
	

	2.3
	The crank handles shall be foldable or recessed to avoid injury.
	
	

	2.4
	It shall be constructed with no sharp edge, and child-safe design.
	
	

	2.5
	The safe working load shall be at least 45 kg.
	
	

	2.6
	Dimensions
	
	

	2.6.1
	The overall length (including revolving bumpers) shall be ≤1850mm.
	
	

	2.6.2
	In-bed length (excluding head and foot-rails, revolving bumpers) shall be within the range of 1470-1850mm.
	
	

	2.6.3
	The overall width (including side-rails, revolving bumpers) shall be ≤ 960mm.
	
	

	2.6.4
	In-bed width (excluding side-rails, revolving bumpers) shall be at least 750mm.
	
	

	2.6.5
	The overall height of the bed shall be within 1500-1600mm.
	
	

	2.6.6
	The height of the bed (measured from the floor to the top of the mattress platform) shall be at least 600 mm.
	
	

	2.6.7
	The highest position of the side rails (all the way up) shall be at least 800mm from the mattress platform.
	
	

	2.7
	Platform Adjustment
	
	

	2.7.1
	The backrest shall be able to raise at least 0-50 degree.
	
	

	2.8
	Construction
	
	

	2.8.1
	The manufacturers work shall use a certified quality management system.  It shall conform to EN13485, which cover the production of medical equipment or equivalent. Documentary proof for compliance shall be submitted. 
	
	

	2.8.2
	The mattress platform shall be flat and solid with smooth-finishing.
	
	

	2.8.3
	The frame of the Paed Bed shall be made of durable and light-weight metals with corrosion, rust and scratch resistant properties/coating.
	
	

	2.8.4
	There shall be 4 revolving non-staining bumpers at the 4 corners of the bed.
	
	

	2.8.5
	The bumper diameter shall be more than 55mm.
	
	

	2.8.6
	The side rails shall be collapsible, lockable, with child-proof mechanism.
	
	

	2.8.7
	The side rails shall be adjustable, lockable at different height.
	
	

	2.8.8
	The side rail locking mechanism shall require two separate, distinct, and simultaneous actions to unlock, preventing accidental or independent operation by the paediatric patient.
	
	

	2.8.9
	The side rail on each side of the bed can be lowered to at least to the level of the bed mattress or be removed to provide unobstructed access.
	
	

	2.8.10
	If the headboard and footboard are removable, there shall be auto-locks when the head and footboard are being placed back onto the bedframe to prevent injuries.
	
	

	2.9
	Castors
	
	

	2.9.1
	There shall be 4 swivelling castors with diameter at least 95mm.
	
	

	2.9.2
	Castors can be locked individually by braking pedal or locked simultaneously by a central braking pedal or an equivalent device.
	
	

	2.9.3
	There shall be at least one anti-static castor to prevent electrostatic discharge.
	
	

	3.
	Item 2.2: General Mattress (“Mattress”)
	
	

	3.1
	The mattress shall be a BS 7177 compliant foam mattress, at least 80mm thick, with a waterproof cover.
	
	

	4.
	Item 2.3 Intravenous Drip Pole (“IV Drip Pole”)
	
	

	4.1
	A one-hand operated telescopic IV Drip Pole made of anti-corrosive materials with at least two (2) open mouth U-shaped or C-shaped hooks.
	
	

	4.2
	The safe working load of the hook shall be at least 1 kg each.
	
	

	4.3
	The height of the IV Drip Pole shall be telescopically adjustable and capable of being fixed at multiple height levels. Functionally equivalent design is also acceptable.
	
	

	4.4
	The telescopic movement shall be adjustable without excessive force and trapping the operation’s fingers. The handheld component for telescopic adjustment shall be non-slippery for firm grasping.
	
	

	4.5
	The IV Drip Pole shall fit securely into the IV drip receptacles that are welded (or locked) at the head or foot corners of the Bed without loose movement and without using any tools.
	
	

	4.6
	The IV Drip Pole shall be located without any obstacle to the lateral side of the Bed to ensure clear pathway for patient transfer.
	
	

	4.7
	The IV Drip Pole shall be located without tilting forward during powered movement of the Paed Bed head section.
	
	

	5.
	Item 2.4 Oxygen Cylinder Holder (“Carrier”)
	
	

	5.1
	The Carrier shall be made of anti-corrosive materials with vertical or horizontal design.
	
	

	5.2
	The Carrier shall be secure in place with no loose movement when oxygen cylinder is placed.
	
	

	(C)
	Training
The Contractor shall provide Training meeting the following requirements:
	
	

	1.
	At least two (2) on-site comprehensive equipment operation and maintenance training courses shall be provided to the CMHHK Operator’s operational personnel and maintenance staff within two (2) weeks after completion of Acceptance Test.
	
	

	2.
	The training courses shall enable CMHHK Operator’s operational personnel and maintenance staff acquires knowledge on day-to-day operation, safety knowledge, routine maintenance and fault diagnosis.
	
	

	3.
	The training shall be conducted in Cantonese with English terms.  All training and training materials provided shall be in Traditional Chinese or English
	
	

	(D)
	Documentation
The Contractor shall provide operation and maintenance (“O&M”) manuals including requirements as below
	
	

	1.
	The Contractor shall submit at least two (2) sets of the manufacturer’s O&M manuals in English or in Traditional Chinese/Simplified Chinese as directed by the Government within two (2) weeks after completion of Acceptance Test. The Contractor shall submit the documentation in form of CD/DVD.
	
	

	2.
	The content of the O&M manuals shall include, but not limited to the following information under separate sections where applicable:
(a)	(a) description of the Goods
(b)	(b) spare parts and consumables, and special tools list
(c)	(c) manufacturers’ certificates
(d)	(d) safety precautions for operation and maintenance
(e)	(e)operation instructions
(f)	(f) maintenance instructions
(g)	(g) maintenance schedules; and
(h)	(h) drawing lists and drawings
	
	

	(E)
	Acceptance Test
	
	

	1.
	On the date(s) specified in the Acceptance Test Plan, the Contractor shall demonstrate, in the presence of the Government Representative, that the Goods is properly tested by the Contractor; in compliance with all mandatory features as stated in the Technical Specifications. 
	
	

	2.
	The Contractor shall conduct a step-by-step testing of each mandatory features as stated in the Technical Specifications. The Government Representative may require the Contractor to make any amendments or adjustments during the tests or request to conduct further reasonable tests.  The Contractor shall be responsible for the provision of all the test equipment and test gear, without generated any additional cost to the Government and the CMHHK Operator.
	
	

	3.
	All tests shall be witnessed by the Government Representative and a minimum of seven (7) days written notice shall be given to the Government Representative before commencement of testing.
	
	

	4.
	The Contractor shall supply, at his own expense, all other necessary materials, as well as other necessary equipment and materials sufficient to complete the testing including repeated tests to the satisfaction of the Government Representative.  
	
	

	5.
	Unless otherwise indicated, the Contractor shall include the cost of all tests, necessary instrument, plant, supervision and labour both at works and on site. All instrument must be properly calibrated, and the accuracy of the test instrument shall be demonstrated where so directed by the Government Representative.  The Contractor shall also carry out all such additional tests as the Government Representative may require to establish that the Goods meets the requirements of the Technical Specifications.
	
	

	6.
	Any defects of workmanship, materials and performance, mal-adjustments or other irregularities which become apparent during the tests shall be rectified by the Contractor at his own expense and the tests shall be repeated at the Contractor's expense to the satisfaction and acceptance of the Government Representative.
	
	

	7.
	All tests, including those to comply with statutory requirements, shall be properly certified in a manner to be agreed with the Government Representative and triplicate copies of all certificates shall be submitted to the Government Representative and appropriate authorities on completion.
	
	

	8.
	The Goods shall be subject to functional and safety testing for their conformance with the operational and reliability requirements.  The tests shall comply with requirements of IEC 60601-2-52 or equivalent.
	
	

	9.
	Only when the Goods has passed all the tests and so certified and all test figures and other relevant information have been recorded in the prescribed manner and accepted by the Government Representative shall the Goods be considered fit for handing over to the Government Representative and the CMHHK Operator.
	
	

	(F)
	Indicative Warranty Service
	
	

	1.
	The Contractor shall provide information to the Government in writing within 24 hours on (a) faults and/or deficiencies known to the Contractor found in Goods supplied by the Contractor which are similar to the Government’s Goods and (b) methods developed to deal with such faults, and/or deficiencies.
	
	

	2.
	Where the same fault or deficiency occurs more than one time within three (3) months, the Contractor shall provide assistance to the Government and the CMHHK Operator in finding the cause(s) and the remedies of such fault or deficiency and preventing further occurrences.
	
	

	3.
	The Contractor shall conduct safety investigation and provide investigation report with safety recommendations when there is a safety incident happened within the Warranty Period of the Goods.
	
	

	4
	Preventive Maintenance
	
	

	4.1
	The Contractor shall provide preventive maintenance services according to manufacturer’s instructions / recommendations, or at least one (1) time per year if not specified, within the Warranty Period, covering full inspection, lubrication and testing in order to keep the Goods in good operating condition and in accordance with the Technical Specifications. The scope of maintenance shall include the following activities: 

(i)	checking and performing safety tests for ensuring satisfactory performance of the Goods;
(ii)	checking and performing all necessary repairs and adjustments of the Goods; and
(iii)	lubricating all components and replacement as necessary, to maintain the Goods in good order including supply of material and labour.
	
	

	4.2
	One (1) month before the start of the Warranty Period, the Contractor shall submit for the Government’s approval a preventive maintenance services schedule for the Goods.  The scheduled services should normally be carried out during normal business hours where other operation of the CMHHK shall not be interfered. The Contractor shall follow the schedule approved by the Government.
	
	

	4.3
	The Contractor shall perform an inspection to check over the Goods and submit a report to the Government within one (1) month before the end of the Warranty Period. The report shall include present condition of the Goods, any irregularities or faults found with cause of the fault identified, remedial actions taken or to be taken, parts repaired/replaced, and any follow-up actions or recommendations.
	
	

	4.4
	In the event of any irregularities identified during the preventive maintenance, the Contractor shall take remedial actions and report to the Government and the CMHHK Operator immediately.  The Contractor shall, based on the findings from the preventive maintenance, draw the attention of the Government and the CMHHK Operator immediately to take any preventive measures or appropriate action at an early stage in order to avoid major faults or degradation in performance of the Goods.
	
	

	5
	Corrective Maintenance
	
	

	5.1
	The Contractor shall provide a hotline for fault reporting to ensure prompt fault attendance. The faults and complaints response time shall be within two (2) working days from the reporting of fault to the Contractor.
	
	

	5.2
	The Contractor shall rectify faults and perform replacement of parts, adjustments and calibrations as may be necessary to ensure the Goods works properly in accordance with the Technical Specifications.
	
	

	5.3
	In the event that the Contractor decides to replace a component of the Goods, the replaced component shall become the property of the CMHHK Operator.
	
	

	6.
	In case of breakdown of the Goods in the course of operation, or the maintenance of the Goods by the Contractor within the Warranty Period, all associated works, and the possible implications shall be borne by the Contractor. 
	
	

	7.
	During the Warranty Period, the maintenance staff of the Contractor shall complete the site record “Maintenance Log Book” in either English or Chinese after each on-site visit.  The entries shall give a full report of the works undertaken during the attendance, including description of fault, zone/areas affected by the fault, cause of fault, remedial actions taken or to be taken, parts repaired/replaced, any follow-up actions or recommendations.
	
	

	8.
	The Contractor shall inform the Government and the CMHHK Operator immediately of any intention by it or the relevant manufacturers to cease manufacture of any Spare Parts prior to any such cessation.  The Contractor shall, at its own cost, be responsible to keep sufficient stock of Spare Parts, and shall ensure that they are fully functional and in good working conditions.
	
	




Part 4 – Implementation Plan
(Note to Suppliers: Please provide the estimated time periods required for the completion of the following tasks, counting from the date of issue an order (“Order Date”).  Both the start and end date of the Order Date is referenced as Month 0.  The Goods should be Ready for Use in the last month of the Implementation Plan.)

Item 1 - Electric Hospital Bed and Accessories 
	Tasks of the Implementation Plan
	Estimated Time Period for 
Performing the Tasks
(The Order Date is set as Month 0)

	
	Start (Month)
	End (Month)

	(a) 
	Order Date (i.e. the date of order placed by the Government, if any) 
	0
	0

	(b) 
	Delivery of the Goods 
	
	

	(c) 
	Delivery of Documentation (Please refer to section D in Part 3 for details)
	
	

	(d) 
	Training (Please refer to section C in Part 3 for Details)
	
	

	(e) 
	Acceptance Tests
	
	

	(f) 
	Goods Ready for Use 
	
	




Item 2 - Paediatric Ward Bed and Accessories
	Tasks of the Implementation Plan
	Estimated Time Period for 
Performing the Tasks
(The Order Date is set as Month 0)

	
	Start (Month)
	End (Month)

	(a) 
	Order Date (i.e. the date of order placed by the Government, if any) 
	0
	0

	(b) 
	Delivery of the Goods 
	
	

	(c) 
	Delivery of Documentation (Please refer to section D in Part 3 for details)
	
	

	(d) 
	Training (Please refer to section C in Part 3 for Details)
	
	

	(e) 
	Acceptance Tests
	
	

	(f) 
	Goods Ready for Use 
	
	




Part 5 – Information on Compliance with International, National and other Recognised Standards or Certifications (if applicable)
(Note to Suppliers: Please indicate in the box below whether the proposed hospital bed and accessories can meet with the standards stated in Column I by inserting a tick in an appropriate box under Column III. If your proposed hospital bed and accessories does not meet the standards stated in Column I, please indicate the equivalent standards met by your proposed hospital bed and accessories in Column IV.  In any case, please attach copies of relevant valid certificates to prove compliance with such standards.)

Item 1 Electric Hospital Bed and Accessories 
	Column I
	Column II
	Column III
	Column IV

	International, National and other Recognised Standards or Certifications
	Requirements 
	Comply with the Standard in Column I?
	Comply with the following equivalent standard
(If “No” in Column III)

	
	
	Yes
	No
	

	IEC60601-1
	
	
	
	

	IEC 60601-1-2
	
	
	
	

	IEC 60601-2-52
	
	
	
	

	IEC 60227
	
	
	
	

	IEC 60245
	
	
	
	

	EN 12531
	
	
	
	

	BS7177
	
	
	
	

	Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify)

	



	


	
	
	



Item 2 - Paediatric Ward Bed and Accessories
	Column I
	Column II
	Column III
	Column IV

	International, National and other Recognised Standards or Certifications
	Requirements 
	Comply with the Standard in Column I?
	Comply with the following equivalent standard 
(If “No” in Column III)

	
	
	Yes
	No
	

	BS7177
	
	
	
	

	IEC 60601-2-52
	
	
	
	

	BS EN 50637
	
	
	
	

	IEC 80601-2-89
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	
	
	
	
	

	Compliance with other international, national and recognised standard(s) or certification(s) in addition to the above (please specify)

	
	




	
	
	




Part 6 – Information on Licencing, Marketing Authorization and MDACS Listing (if applicable)
(Note to Suppliers: Please advise whether your company and the proposed Goods have the following licence, marketing authorization and Medical Device Administrative Control System (“MDACS”) listing. If affirmative, please provide copies of relevant licences, confirmation and certificates for our reference.) 
Item 1 - Electric Hospital Bed and Accessories 
	Question
	Licensing/Certification/Listing Information of the System
	 (Please tick in the appropriate box)

	
	
	#Yes
	No

	1
	Does the proposed Goods have marketing authorization of the European Union (EU) for affixing of CE marking on the product?
	
	

	2
	If the proposed Goods has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable).
(a) *Declaration of conformity by the manufacturer; or
(b) *Certificate of conformity issued by a notified body.
	
	

	3
	Does the proposed Goods have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”.
Country / Region :  _________________________________
	
	

	4
	Has your proposed Goods been listed in the MDACS of the Department of Health (if applicable)? 
	
	

	5
	What class of medical device is your proposed Goods (if applicable)?
(a) EU : Class ______
(b) United States : Class ______
(c) Other country/region (please specify below):
· Country/Region  _________
· Class _________
	
	

	6
	Does the proposed Goods has marketing authorization of verifying the maximum loading capacity? 
	
	

	7
	Does the proposed Goods has marketing authorization of the flammability standard (e.g. BS 7177:2008+Al: 2011 (Medium Hazard) when tested in accordance with BS EN 597-1:2015 (Ignition source 0), BS EN 597-2:2015 (Ignition source 1) and BS 6807:2006 Section 2 (Ignition source 5)?  
	
	


#Please provide a copy of the licence/confirmation/certificate for reference.





Item 2 - Paediatric Ward Bed and Accessories 
	Question
	Licensing/Certification/Listing Information of the System
	 (Please tick in the appropriate box)

	
	
	#Yes
	No

	1
	Does the proposed Goods have marketing authorization of the European Union (EU) for affixing of CE marking on the product?
	
	

	2
	If the proposed Goods has marketing authorization of EU, please state the type of supporting document (*delete which is not applicable).
(a) *Declaration of conformity by the manufacturer; or
(b) *Certificate of conformity issued by a notified body.
	
	

	3
	Does the proposed Goods have marketing authorization in country/region other than United States and EU? Please specify below if your answer is “Yes”.
Country / Region :  _________________________________
	
	

	4
	Has your proposed Goods been listed in the MDACS of the Department of Health (if applicable)? 
	
	

	5
	What class of medical device is your proposed Goods (if applicable)?
(d) EU : Class ______
(e) United States : Class ______
(f) Other country/region (please specify below):
· Country/Region  _________
· Class _________
	
	

	6
	Does the proposed Goods has marketing authorization of verifying the maximum loading capacity? 
	
	

	7
	Does the proposed Goods has marketing authorization of the flammability standard (e.g. BS 7177:2008+Al: 2011 (Medium Hazard) when tested in accordance with BS EN 597-1:2015 (Ignition source 0), BS EN 597-2:2015 (Ignition source 1) and BS 6807:2006 Section 2 (Ignition source 5)?

	
	


#Please provide a copy of the licence/confirmation/certificate for reference.



Part 7 – Indicative Price Information
(Note to Suppliers: The price information provided in this Part 7 is for Government’s consideration only and shall not constitute any commitment on the part of the Government or your company. Nevertheless, please provide the information as accurate as possible.)

Indicative Price Information for the Goods
Item 1 - Electric Hospital Bed and Accessories
	Item
	Description
	Estimated
Quantity

	Unit Price
	Estimated Goods Price

	
	
	
	One-time Unit Price
(HK$)
	Estimated Goods Price for the Item specified opposite
(HK$)

	
	
	(a)
	(b)
	(c) = (a) x (b)

	1
	Supply and delivery of the following Goods (electric hospital bed, general mattress, lifting pole, IV drip pole and oxygen carrier) including the provision of a minimum 12-months warranty period. 
	
	
	

	1.1
	Electric Hospital Bed
	213
	
	

	  1.2
	General Mattress
	213
	
	

	1.3
	Lifting Pole
	213
	
	

	1.4
	Intravenous Drip Pole 
	213
	
	

	1.5
	Oxygen Cylinder Carrier 
	11
	
	

	  1.6
	Provision of training services as detailed in section C in Part 3  
	2 courses
	
	

	  1.7
	Documentation as detailed in section D in Part 3
	1 lot
	
	

	  1.8
	Other (please specify)
	(please specify)
	
	

	Total One-time Charge
(i.e. Sum of Estimated Goods Prices of Item 1.1-1.8)
	


Note: * The Total One-time Charge shall include one-year of warranty period.


Item 2 - Paediatric Ward Bed and Accessories 
	Item
	Description
	Estimated
Quantity

	Unit Price
	Estimated Goods Price

	
	
	
	One-time Unit Price
(HK$)
	Estimated Goods Price for the Item specified opposite
(HK$)

	
	
	(a)
	(b)
	(c) = (a) x (b)

	2
	Supply and delivery of the following Goods (paediatric ward bed, general mattress, IV drip pole and oxygen carrier) including the provision of a minimum 12-months warranty period. 
	
	
	

	2.1
	Paediatric Ward Bed 
	8
	
	

	2.2
	General Mattress
	8
	
	

	2.3
	Intravenous Drip Pole 
	8
	
	

	2.4
	Oxygen Cylinder Carrier 
	2
	
	

	 2.5
	Provision of training services as detailed in section C in Part 3  
	2 courses
	
	

	 2.6
	Documentation as detailed in section D in Part 3
	1 lot
	
	

	  2.7
	Other (please specify)
	(please specify)
	
	

	Total One-time Charge
(i.e. Sum of Estimated Goods Prices of Item 2.1-2.7)
	


Note: * The Total One-time Charge shall include one-year of warranty period.


Part 8 – Indicative Maintenance Charges and Spare Parts Price

(Notes to Suppliers for completion of Part 8)
(1) Pursant to Part 7 above, the proposed Goods shall have a warranty period of not less than 12 months. The indicative warranty service requirements are stipulated in section F in Part 3, which are subject to changes at the sole discretion of the Government.
(2) It is expected that the maintenance services shall be comprehensive, all inclusive and shall cover all parts, components, labour and software support services. If your company considers that any components of the Goods may not be covered by the maintenance services (saving that the labour shall always be covered by the maintenance services) and may need to be charged separately, please indicate replacement costs of these components and their replacement frequency.
(3) The annual maintenance charge within the serviceable life of the proposed Goods is adjustable in accordance with the consumer price index (B) upon the expiry of each 12-months period of maintenance service. 

Item 1 - Electric Hospital Bed and Accessories
(a) Indicative Maintenance Prices of the Goods
	Year
	Annual Maintenance Charge
(HK$ per annum)

	First 12-months period of maintenance service after the end of warranty period
	



(b) Indicative Replacement Prices of Equipment’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)
(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

	Item
	Name of Items
	Indicative 
Replacement Price (HK$/no.)
	Indicative Replacement Frequency (e.g. once every 3 years)

	1
	
	
	

	2
	
	
	





(c) Indicative overtime charges for provision of maintenance services after office hours (if applicable)
(Office hours mean 9 am to 5 pm from Monday to Friday excluding public holidays)  

	(a)
	Rates of overtime charges for maintenance service outside the office hours
	HK$                   per hour

	(b)
	Minimum service hour(s) per call
	            service hour(s) per call




(d) Indicative Prices for Replacement of Other Spare Parts (if applicable)

	Item
	Name of Items
	Price (HK$/no.)
	Indicative Replacement Frequency (e.g. once every 3 years)
	Expected time for delivery 
(weeks)

	1
	
	
	
	

	2
	
	
	
	




Item 2 - Pediatric Ward Bed and Accessories 
(a) Indicative Maintenance Prices of the Goods
	Year
	Annual Maintenance Charge
(HK$ per annum)

	First 12-months period of maintenance service after the end of warranty period
	



(b) Indicative Replacement Prices of Equipment’s Components not covered by the Maintenance Services (if applicable) (Leave the following table blank if not applicable)
(Note to Suppliers: The labor costs for replacement of these components shall always be covered by the maintenance charges for the provision of the maintenance services regardless whether the prices for the supply of these components are covered by the maintenance services or not.)

	Item
	Name of Items
	Indicative 
Replacement Price (HK$/no.)
	Indicative Replacement Frequency (e.g. once every 3 years)

	1
	
	
	

	2
	
	
	



(c) Indicative overtime charges for provision of maintenance services after office hours (if applicable)
(Office hours mean 9 am to 5 pm from Monday to Friday excluding public holidays)  

	(a)
	Rates of overtime charges for maintenance service outside the office hours
	HK$                   per hour

	(b)
	Minimum service hour(s) per call
	            service hour(s) per call




(d) Indicative Prices for Replacement of Other Spare Parts (if applicable)

	Item
	Name of Items
	Price (HK$/no.)
	Indicative Replacement Frequency (e.g. once every 3 years)
	Expected time for delivery 
(weeks)

	1
	
	
	
	

	2
	
	
	
	





Part 9 – Supplementary Information

A. Sales Volume of the Offered Goods (leave blank if information is not available)

	Description
	Annual Sales for the past three years

	Item 1 - Electric Hospital Bed and Accessories
	In Hong Kong:          sets

	
	Globally:               sets

	Item 2 - Paediatric Ward Bed and Accessories 

	In Hong Kong:          sets

	
	Globally:               sets











Part 10 – Questionnaires

Item 1 - Electric Hospital Bed and Accessories
	Information Required
	Complete by Suppliers
(use separate sheet, if needed)

	(1) Any details on parts and services covered in Warranty Service in addition to Part F?
	

	(2) Would a 2-year contract period (starting from the date specified in letter of acceptance) acceptable to your company? Order will be placed by 2 to 3 batches within the contract period.
	

	(3) What is the payment schedule?

	



Item 2 - Paediatric Ward Bed and Accessories
	Information Required
	Complete by Suppliers
(use separate sheet, if needed)

	(1) Any details on parts and services covered in Warranty Service in addition to Part F?
	

	(2) Would a 2-year contract period (starting from the date specified in letter of acceptance) acceptable to your company? Order will be placed by 2 to 3 batches within the contract period.
	

	(3) What is the payment schedule?



	



END
